## Slide 1
Design and tomography test of soft X-ray array diagnostics on KSTAR
Seung Hun Lee*, Juhyeok Jang, Joohwan Hong, H. Y. Lee,
Siwon Jang, and Wonho Choe**
Korea Advanced Institute of Science and Technology (KAIST), Daejeon, Korea
Fusion Plasma Transport Research Center (FPTRC), Daejeon, Korea
*leesh81@kaist.ac.kr **wchoe@kaist.ac.kr
PPPL, Feb. 17. 2014

## Slide 2: Outline
KSTAR status & Motivation
Signal estimation & Tomography test
  Absolute calibration of AXUV
  KSTAR-like SXR images L- and H-modes
  Photocurrent of the detector with various filter sets
  Various reconstruction methods
  Reconstruction results of KSTAR-like phantoms with a P-T method
Available research fields
  Impurity transport (Ar, W)
  Sawtooth crash
  MHD mode identification
Engineering design
  Detector head & Installation scheme
  2-D map of LOS
  DAQ system
Discussion

## Slide 3: Daejeon - Korea
East sea
West sea
Daejeon

## Slide 4: KAIST & NFRI

## Slide 5: KSTAR (Korea Superconducting Tokamak Advanced Research)
Superconducting Tokamak
Complete construction : Sep 14, 2007
Operation data : 2008 –
Major radius : 1.8m
Minor radius : 0.5m
Plasma volume : 17.8m3
Plasma current : >2.0MA (1MA)
Toroidal field : >3.5T (2T)
Pulse length : >300s (10s)
Plasma fuel : H, D-D
Auxiliary heating : ~ 28MW

## Slide 6: KSTAR for 2011 Campaign
W. C. Kim, et al., “Overview of the KSTAR 2011 experimental results”, 3th KSTAR conference, 2012

## Slide 7: Map of KSTAR diagnostics
XICS #1 (’09 : normal)
X-ray camera (MWPC ’11)
(u&l) Bolometer-divertor
Recip. Langmuir probe
Fast-ion detector
Tangential X-ray camera (’11)
(t) Recip. Langmuir probe
Soft X-ray array (’09)
TV #3 (’10 : 1kfps)
IRTV #1 (’10)
Bolometer-imaging (’10)
Ellipsometry
mm-Interferometer (’12)
Reflectometer (’12~)
X-ray Pulse Height Anal.
Hall probe (’08)
IRTV (normal)
XICS #2 (’10 : tangential)
X-ray camera (GEM, ’12)
FIR Interferometer (’12 : 5ch)
VUV Survey Spec. (’12)
Edge SXRA (’14)
mm-Interferometer (’08~’11)
Reflectometer (’09~’11)
ECEI (’12)
MIR (’14)
(t&b) FIR (’11 : 1ch)
ECEI (’10~’11)
Illuminator (’08)
TV #2 (’08)
CES-background (8/50ch)
IRTV #2 (’11:tangential)
TS - Optics (’10 : 5/42ch)
(u&l) IRTV (divertor)
CES (’10 : 8/50ch)
Coherence Imaging (’11)
MSE (’12)
BES (’12)
Bolometer - resist (’09) (12ch)
TS - Laser (’10)(2J 10Hz)
CX-NPA (’12)
(t&b) Laser for D. Thomson
(t&b) Laser for LIF
ECE (’08)
ECE Grating Polychrometer
Diagnostic Neutral Beam
H-alpha Monitor (’08)
Visible Spectrometer (’08)
Visible Filterscope (’08)
Visible Brems. Array
Multi-chord Vis. Spec.
(u&l) CES-poloidal
(l) Recip. Langmuir Probe
Inspection Illuminator
TV #1 (’08)
IRTV #3 (’12 : tangential)
MSE(II)
LHCD
ICRF
Pellet Injector
NBI 1
ECCD
NBI 2
Pumping Duct
IVCC Leads
IVCC Leads
ECH
IVCC Leads
IVCC Leads
Installed
Planned from 2014
To be decided
Fuel
GDC
GDC
Hard X-ray (’09)
Neutron detector (’10)
A
B
C
D
E
F
G
H
I
J
K
L
M
N
O
P

## Slide 8: Motivation
Demand
  A two-array system has a limitation to reconstruct the plasma with elongation  and triangularity  with high accuracy.
  At least 6 arrays are required to describe  and  .
In 2014, a four-array system of soft X-ray diagnostics is under preparation on KSTAR.
Phantom
6 arrays (TBD)
4 arrays (2014)
2 arrays (2009)
Elongation
Triangularity

## Slide 9: Coverage test for array system
HL
HC
HU
VU
VL
HU
HL
VL1
VU1
VL2
HL
VU
HU
3 arrays
5 arrays
6 arrays
VU2
HU
HL
VU
VL
4 arrays

## Slide 10: Calculation for signal estimation & Tomography reconstruction

## Slide 11: Signal estimation
Properties of detector system
Absolute calibration
Bremsstrahlung calculation in KSTAR L- and H-modes
Incident photon rate & Photocurrent of the detector
Filter sets for Te measurement
  Be filters (10, 50, 250, 400 μm)

## Slide 12: Photodiode detector & Pre AMP
AXUV-16ELG detector
53 mm
15 mm
Requirement
  Fast response ~ MHz
  High sensitivity to XUV and soft X-ray
Specification
  Active area: 5  2 mm2
  Shunt resistance: 100 m
  Capacitance: 2 nF
  Rise time (10-90%): 0.5 μs
  Gain: 106 V/A
  Detection efficiency: 0.27 A/W*
AXUV-16ELG array
AMP-16 remote panel
AMP-16 main circuit
Ribbon cable
55 mm
73 mm
*See p.10
t = 2 s
r = 2 cm

## Slide 13: Solid angle
Plasma volume, dVp
h
Aperture, Aap,i
Detector, Adet,i
di
Line of sight, Li
Thickness, dli
2.5 × 1 mm2
5 × 2 mm2
92 mm
dPi : measured power emitted from the plasma volume dVp
ci : calibration factor

## Slide 14: Absolute calibration
Be window (0.25 mm)
Be window (0.254 mm)
Si p-i-n diode
(7 mm2  0.3 mm)
Al filter
Distance
Fe-55
MOXTEK
X-ray source
Slit
Experimental setup
Spectrum of the source
Response vs Itube
Response vs Vtube
Response vs Filter
Detection efficiency of AXUV
Collaboration with ENEA-Frascati

## Slide 15: Typical ne & Te in KSTAR L- & H-mode
Typical ne & Te profiles of L-mode and H-mode are prepared to calculate X-ray radiation spectrum.
ne profile
Te profile

## Slide 16: ne, Te based SXR profile calculation (after filter)
X-ray continuum radiation and filtered radiation spectra by Be foils were calculated in each case.
L-mode
H-mode

## Slide 17: Synthetic SXR image based on typical KSTAR L-mode
Continuum
Be 10 μm
Be 50 μm
Be 250 μm
Be 400 μm
 = 1.8
 = 0.2
Δmax = 0.1 m

## Slide 18: Synthetic SXR image based on typical KSTAR H-mode
Continuum
Be 10 μm
Be 50 μm
Be 250 μm
Be 400 μm
 = 1.8
 = 0.2
Δmax = 0.1 m

## Slide 19: Line of Sight (LOS)
Reconstruction region
VU
HU
HL
VL
Spatial resolution: 2 cm
32 ch × 4 arrays = 128 ch
Total 128 chords
Coverage test
1
32
1
32

## Slide 20: Expected photo-current of AXUV
Photo-current level is ~0.03μA in L-mode and 0.15 μA in H-mode, which means the voltage signal is ~30 mV in L-mode and 150 mV in H-mode.
Vertical
Horizontal
L-mode
H-mode
Vertical
Horizontal

## Slide 21: Te from multi-energy SXRA
Plasma
SXR
(ne, Te, Zeff)
Thin
filter
Thick
filter
Signal 2
depends only on
Te
Signal 1
Continuum radiation
Assumption
  Maxwellian distribution of electron energy
  Continuum radiation only
Te can be derived from signal ratio.

## Slide 22: Four Be filter sets
Be filters of 10 μm, 50 μm, 250 μm, and 400 μm
Detection efficiency (D) is same with absorption probability of 400 μm Si (AXUV).
Each photon spectra arriving at the detector can be calculated.
D, Prad, and transparency of Be filters
X-ray spectra after Be filter (Te = 2 keV)

## Slide 23: Signal ratio & Dynamic range of the detector
10 & 50 μm
10 & 250 μm
10 & 400 μm
50 & 250 μm
50 & 400 μm
250 & 400 μm
0.4-2.0 keV
1.2-6.0 keV
1.5-9.0 keV
0.7-3.5 keV
1.0-5.5 keV
0.2-1.0 keV
A set of 50 μm and 250 μm is best for KSTAR plasma.

## Slide 24: Tomography test
Reconstruction methods
Prepared phantoms
  (L-mode, H-mode, m/n = 1/1, m/n = 2/2, and m/n = 3/2)
Reconstruction results
2-D Te reconstruction by ME-SXR
MHD mode analysis

## Slide 25: Tomography
x
y
p
Line of sight
Dashed line : Viewing Cone
Plasma
Detector
r


Tomography :
A non-invasive imaging tool for observing the inner structure of the plasmas
Coordinate system for defining tomography system
Acquiring soft x-ray signal ( f(p, ) )
from many lines of sight
Solving inverse problem
(ill-posed problem) :
Reconstruction of 2-D local
emissivity ( g(r, ) )

## Slide 26: Tomographic reconstruction codes developed at KAIST
Discrete method
J0
J1
J2
J3
J4
m = 0
m = 1
m = 2
Bessel function
Zernike polynomial
Cormack
Classified by objective functional O(g)
Minimum norm: Phillips-Tikhonov
Entropy: Maximum entropy
Multi-order smoothness: Minimum Fischer Information
Analytic method

## Slide 27: 6 array reconstruction test
Single peak
Hollow
Method | Relative error,  (%) |  | Cal time (s)*
 | Single | Hollow | 
P-T | 0.59 (b) | 1.55 (e) | 20
MEM | 6.90 (c) | 5.72 (f) | 15
P-T
MEM
*under Pentium 2.67GHz, Windows platform
6 array system
P-T performance is better than conventional MEM in accuracy.

## Slide 28: 4 array reconstruction test (P-T)
L-mode phantom
Reconstruction
Continuum
Be 10 μm
Be 50 μm
Be 250 μm
Be 400 μm

## Slide 29: 4 array reconstruction test (P-T)
L-mode phantom
Reconstruction
Continuum
Be 10 μm
Be 50 μm
Be 250 μm
Be 400 μm

## Slide 30: 2-D Te reconstruction by ME-SXR
Be 10/50 μm
Be 10/250 μm
Be 10/400 μm
Be 50/250 μm
Be 50/400 μm
Be 250/400 μm
Good agreement of Be 10/50 μm in low temperature
It contains reconstruction error.
L-mode

## Slide 31: 2-D Te reconstruction by ME-SXR
Be 10/50 μm
Be 10/250 μm
Be 10/400 μm
Be 50/250 μm
Be 50/400 μm
Be 250/400 μm
Good agreement in r/a < 0.6
It contains reconstruction error.
H-mode

## Slide 32: Phantoms with MHD modes
m/n = 1/1
@ r/a = 0.3
m/n = 2/2
@ r/a = 0.3
m/n = 3/2
@ r/a = 0.5

## Slide 33: Reconstruction results (P-T)
m/n = 1/1
@ r/a = 0.3
m/n = 2/2
@ r/a = 0.3
m/n = 3/2
@ r/a = 0.3

## Slide 34: MHD mode identification by SVD analysis (m/n = 1/1)
MHD m/n = 1/1 mode is clearly seen by the SVD method.
Two phases are distinguishable with /2 difference.
k = 1
k = 2
k = 3

## Slide 35: MHD mode identification by SVD analysis (m/n = 2/2)
MHD m/n = 2/2 mode can be analyzed by the SVD method.
Two phase of /2 can be distinguished.
k = 1
k = 2
k = 3

## Slide 36: MHD mode identification by SVD analysis (m/n = 3/2)
MHD m/n = 3/2 mode can be analyzed by the SVD method.
Two phase of /2 can be distinguished.
k = 1
k = 2
k = 3

## Slide 37: Expected studies

## Slide 38: Impurity transport research (Ar, W)
50 m
W cm-2 eV-1
Two filter sets can distinguish the different species of impurities.
SANCO can provide transport coefficients of impurities
Ar
W
Continuum radiation
KSTAR X-ray spectrum by calculation (Te = 2 keV, ne = 1019 m-3)
Transport coefficients of impurities by SANCO
J. Hong

## Slide 39: Sawtooth crashes (example)
m = 1 (f = 19 kHz) is shown by spectrogram.
Maximum displacement from the initial position: 0.13 m
Maximum rotation speed: 10.7 km/s
Spectrogram
m = 1
f = 19 kHz
Trajectory of the hot core (rtEFIT)
KSTAR
Notes: 마찬가지로 H-mode에서 결과입니다. 주파수는 19 kHz로 L-mode보다는 더 높았고 최대 속도 10.7 km/s였습니다. 중심의 이동 거리는 0.13 m로 L-mode와 거의 같았습니다.

## Slide 40: Comparison between L- & H-mode
Crash
Up to 5 km/s
Up to 0.1 m
Crash
Up to 10 km/s
Up to 0.1 m
L-mode, low vФ
H-mode, high vФ
Crash in three steps
Crash in one step
Displacement from central position
Displacement from central position
Poloidal velocity
Poloidal velocity
KSTAR
Notes: L-mode와 H-mode 비교를 하자면, crash가 일어나는 양상이 앞서와 같이 다릅니다. 확대한 신호를 보면 L-mode 에서는 crash가 한 번에 일어나 oscillation 피크값이 최대가 되었다가 바로 잠잠해지는 모양이고, H-mode에서는 최대 피크점이 있고 그 이후에도 oscillation이 세 번에 걸쳐 그 진폭이 잦아드는 모양입니다. (9, 11 페이지 그림과 일치)

## Slide 41: Correlation between m=1 frequency & vФ
The m=1 frequencies are compared with toroidal velocity for all the sawtooth crashes.
Rotation frequency can be calculated by
The m=1 frequency is determined by rotation velocity
KSTAR
Notes: 이러한 m=1 모드의 주파수와 토로이달 회전 속도와 비교해보면 이와 같습니다. 6개의 실험에서 일어난 모든 sawtooth crash를 fft하여 알아낸 주파수와 그 시간에 해당하는 토로이달 회전속도와 비교한 값입니다. 이에 따르면 m=1 모드의 주파수는 토로이달 회전 속도에 의해 결정된다고 할 수 있겠습니다.

## Slide 42: MHD mode analysis with SVD
Time sequences of reconstructed emissivity
Spatial structures
Temporal evolution
Singular Value
k = 1
k = 2
k =3
k = 4
k = 5
k = 6
k =7
k = 8
No m=2
KSTAR
Notes: 이것은 SVD로 모드 알아내는 그림입니다. m=2 모드는 보시다 시피 나타나지 않았습니다.

## Slide 43: 4 array SXRA preparation
Detector head
  AXUV16-ELG
  Filter wheel
Installation scheme
  Vertical & Horizontal arrays
2-D map of LOS

## Slide 44: Detector Head
AXUV-16ELG detector
53 mm
15 mm
Specification
  Active area: 5  2 mm2
  Shunt resistance: 100 m
  Capacitance: 2 nF
  Rise time( 10-90%): 0.5 μs
  Preamp gain: 106 V/A
  Bandwidth: 600 kHz
AXUV
Preamp
Cover
Viewing cone
Filter wheel
Different filter sets
Four AXUV detectors in a detector head
  Two for thin Be filter, and the others for thick Be filter
Preamp (106 V/A) close to the detectors
Filter wheel can provide other filter sets.
  (e.g.) Ar Ross filter, Bolometer mode
  Not to be used in 2014
t = 2 s
r = 2 cm

## Slide 45: Position of Vertical & Horizontal arrays
Reconstruction area
Vertical array
Horizontal array
Passive stabilizer
Vertical array should be attached behind the passive stabilizer.
Horizontal array will be mounted on vessel wall (exit of a cassette).

## Slide 46: Installation Scheme (Vertical)
Through-hole (ϕ65) on passive stabilizer for viewing cone
Insulation between P/S and Detector head is needed.
Detector head
Signal outlet
Passive stabilizer
Insulator
Flattened
ϕ65 Hole
M6 × 6 Bolts

## Slide 47: Specific measurements (Vertical)
Distance between the detector head and IVCC is 50 mm.
Mechanical vibration due to the disruption
IVCC
Passive stabilizer
C tile
IVCC
Detector head
M6 × 6
50 mm
130 mm
202 mm
Size in mm

## Slide 48: Installation Scheme (Horizontal)
Installation on the vessel wall with tapped holes (M4 × 12)
Horizontal array will be mounted on vessel wall (exit of a cassette).
Detector head
Apertures
Mounting pad
Vessel wall
Cassette

## Slide 49: Line of Sight (LOS)
Reconstruction region
VU
HU
HL
VL
Spatial resolution: 2 cm
32 ch × 4 arrays = 128 ch
Each chord has two different filter sets (e.g. 50 μm & 250 μm).
Total 256 channels
Coverage test

## Slide 50: Data flow of DAQ system
Preamp (106 V/A)
Photo diode
Coaxial cable in differential input
BNC-2090A
SHC68-68-EPM
LTU + carrier
Chassis
Data (256 ch, 500 kS/s)
MDS+ Server
Power supply
D-tAcq
ACQ196CPCI 96 ch × 4
Main Control
Optical signal
Trigger  8
Controller (Linux)
Ethernet
Torus hall
Diagnostic room
Data (256 ch, 500 kS/s)
Controlled by RS232
Optical cable

## Slide 51: Location of Diagnostic devices
A
B
C
D
E
F
G
H
I
J
K
L
M
N
O
P
Entrance
19” Rack
Detector & Preamp
Coaxial Cable 15 m
PXIe Chassis (NI DAQ)
Power Supply
Controller (Linux)
D-tAcq (SXRA)
Optical Cable 100 m
Diagnostic Room
19” Rack
Controller (Linux)
Main Room

## Slide 52: Software error by neutron radiation
http://gigglehd.com/zbxe/newsreport/5713195
2011 International Reliability Physics Symposium

## Slide 53: DAQ error (2012-13 KSTAR campaign)
NI-6123 (500 kS/s, 8 ch)
3 of 8 boards are failed in digital I/O.
Repairing charge: $700/eh

## Slide 54: Discussion
Signal level of SXRA
Idet ~ 0.02-0.12 μA, V ~ 20-120 mV in L-mode
Idet ~ 0.05-0.45 μA, V ~ 50-450 mV in H-mode
Choice of Be filter sets for Te measure
(10 μm, 50 μm, 250 μm, 400 μm, etc.)
Tomography reconstruction quality with 4 array system
Capability of MHD mode identification (m/n = 1/1, 2/2, 3/2)
Installation of in-vessel SXRA diagnostics
Neutron damage to DAQ system

## Slide 55: Supplementary

## Slide 56: Cormack method (analytic)
J0
J1
J2
J3
J4
Decomposed by harmonics
Expansion by basis function(Bessel function)
Calculating f components of each coefficient
Measured signals
Coefficient
Converting to linear equations

## Slide 57: Phillips-Tikhonov (P-T) method
Set
regularization parameter
error
number of detectors
Laplacian of g
minimize J by
Solution is
To find out optimized
Goodness byGCV
should be minimized by

## Slide 58: Iterative solution
Single peak
Hollow
3 iteration
Iteration Number
Relative Error
6 iteration

## Slide 59: Modified P-T method
P-T method is modified for withstanding to the noisy data by nonlinearising the weight matrix
The difference between measured signal from solution g is defined by
Talyor expansion
The step size solution
The kth iterative solution
Jacobian matrix of D
step size of g, h << g
0
Lee et al., Curr. Appl. Phys. 893, 10 (2010)

## Slide 60: Maximum entropy method
Posterior probability density function
Bayesian rule
Prior and Likelihood probability density function
Lagrange Multiplier method
Iterative calculation until the maximum entropy of reconstruction result meets

## Slide 61: Minimum Fisher information
f i = signal from the ith detector
gjτ = local emissivity of the jth pixel at the τth time slice
T = geometrical weight matrix
B = gradient matrix
λ = regularization parameter
S = total number of time slices
L = number of detectors
M = transfer matrix
U = functional matrix
P = number of pixels
wk = inverse of the kth image pixel
σi = standard deviation of the ith measurement
The regularization parameter is iterated until .
Fisher Information
chi-square
Minimizing the functional matrix U 
Minimum Fisher Information method
M = U-1T

## Slide 62: MHD mode in TCV
Identification of poloidal modes using SVD
Tomogaphic
Plane
Tangential 2D array
Tangential
Tomography
Relative Intensity
Toroidal direction
Integrated relative intensity
Identification of m/n using poloidal array combined with tangential 2-D array
Toroidal mode (n = 1) from the emissivity along the red line
Ray-tracing with a rectangular 3-D weight matrix
m=2
m=1

## Slide 63: Data acquisition system
Preamp (105 V/A)
Photo Diode
Coaxial Cable in Differential Input
BNC-2090A
SHC68-68-EPM
LTU + carrier
PXIe-1065
Data (64 ch, 10 kS/s)
MDS+ Server
Power Supply
Data (64 ch, 500 kS/s)
PXI-6123  8
PXIe-8372
Main Control
Controlled by RS232
Optical Signal
Trigger  8
Controller (Linux)
Ethernet