## Slide 1
AIMS briefing to NSTX-U group: Accelerator-based surface diagnostic for plasma-wall interactions science
Dennis Whyte Zach Hartwig, Harold Barnard, Brandon Sorbom, Pete Stahle, Richard Lanza
D. Terry, J. Irby, G. Dekow & C-Mod engineering staff
Plasma Science and Fusion Center
Massachusetts Institute of Technology
CAARI 2012
Fort Worth, TX
08 AUG 12
@
08 AUG 12

## Slide 2: New in-situ diagnostics are required to significantly advance PWI science in magnetic fusion devices
Ex-situ diagnostics and “benchtop” PWI experiments are critically limited
  Unable to replicate tokamak-relevant PWI conditions (“benchtop”)
  Limited PFC surfaces available for measurement (IBA)
  “Archaeological” measurements lack dynamic PWI information (IBA)
In-situ PWI surface diagnostics are severely limited in deployment and unable to meet all requirements
The ideal PFC surface diagnostic would provide measurements:
  in-situ without vacuum break
  on a shot-to-shot frequency for time resolution and PWI dynamics
  of large areas of PFC surfaces (poloidally and toroidally resolved)
  of elemental/isotope discrimination to depths of ~10 microns

## Slide 3: AIMS exploits intra-pulse capabilities of a tokamak and deuteron-induced reactions to investigate PWI
AIMS (Accelerator-based In-situ Materials Surveillance) derives from two key observations:
  The tokamak magnetic fields can be used between plasma shots to steer a charged particle beam to PFC surfaces of interest
  The gammas and neutrons produced by low-energy, deuteron-induced nuclear reactions provide a comprehensive diagnostic tool for PWI
PWI Issue | Nuclear reaction | Induced particle energy (MeV)
Fusion fuel retention | 2H(d,n)3He 3H(d,n)4He | En = 2 – 4 MeV En = 17 – 19 MeV
Erosion / redeposition | 6Li(d,p+g)7Li 8Be(d,p+g)9Be | Eg = 0.478 Eg = 0.718
Wall conditioning | 11B(d,p+g)12B 16O(d,p+g)17O | Eg = 0.953, 1.674 Eg = 0.871
Impurity transport | Accessible Low-Z reactions | Eg <= 5 MeV

## Slide 4: Basic principles for AIMS on Alcator C-Mod: in-situ ion beam analysis
RFQ LinAc
beamline
Radio Frequency Quadrupole (RFQ) linear accelerator injects 0.9 MeV D+ beam into vacuum vessel through a radial port

## Slide 5: Basic principles for AIMS on Alcator C-Mod: in-situ ion beam analysis
Radio Frequency Quadrupole (RFQ) linear accelerator injects 0.9 MeV D+ beam into vacuum vessel through a radial port
Tokamak magnetic fields provide steering via the Lorentz force
RFQ LinAc
beamline
Toroidal field provides poloidal steering :
Vertical field provides toroidal steering :

## Slide 6: Basic principles for AIMS on Alcator C-Mod: in-situ ion beam analysis
Radio Frequency Quadrupole (RFQ) linear accelerator injects 0.9 MeV D+ beam into vacuum vessel through a radial port
Tokamak magnetic fields provide steering via the Lorentz force:
D+ induce high Q nuclear reactions with low Z isotopes in PFC surfaces producing ~MeV neutrons and gammas
RFQ LinAc
beamline
γ
n
Toroidal field provides poloidal steering :
Vertical field provides toroidal steering :

## Slide 7: Basic principles for AIMS on Alcator C-Mod: in-situ ion beam analysis
Radio Frequency Quadrupole (RFQ) linear accelerator injects 0.9 MeV D+ beam into vacuum vessel through a radial port
Tokamak magnetic fields provide steering via the Lorentz force:
D+ induce high Q nuclear reactions with low Z isotopes in PFC surfaces producing ~MeV neutrons and gammas
In-vessel detection and energy spectroscopy provides a wealth of information on PFC surface composition and conditions
RFQ LinAc
beamline
Toroidal field provides poloidal steering :
Vertical field provides toroidal steering :
γ
n
Gamma and neutron detectors
Reentrant
vacuum
tube

## Slide 8: AIMS was fits into the extremely crowded area around C-Mod's horizontal ports
Vacuum vessel
Concrete igloo
(shielding)
HiReX Sr
Spectrometer
(Plasma diagnostic)
Impurity injector
(plasma transport studies)
Hard X-Ray Camera
(fast electron diagnostic)
Lyman-alpha camera
(plasma diagnostic)
Accelerator-based ents
*

## Slide 9: AIMS was designed to fit amidst the extremely crowded area around C-Mod's horizontal ports
A solid model with cutaways of the proposed
installation location for AIMS on Alcator
C-Mod. (Lots of stuff not shown!)

## Slide 10
Installed AIMS: RFQ accelerator, focusing quadrupole beamline, reentrant tube, and particle detectors
Cryopumps
(provide vacuum)
Gate
valves
RFQ cavity
(accelerates
deuterons)
RF input
(provides
accelerating
electric field)
Quadrupole
(beam
focusing)
Neutron and
gamma
detectors
0.9 MeV
deuteron
beam
Reentrant vacuum
tube
Gate valve
(vacuum
barrier)

## Slide 11: AIMS was installed on C-Mod

## Slide 12: Beams can be steered poloidally and toroidally: Will need to explore this with NSTX-U ~cw B capability + ports

## Slide 13: AIMS provides non-perturbing, quantitative measurement of surface isotopes/elements; but is effectively a re-invention of Ion Beam Analysis
Time [10 us/div]

## Slide 14: An 0.9 MeV deuteron RFQ accelerator has been completely refurbished and upgraded
~25 year old prototype RFQ (radiofrequency quadrupole) accelerator from Accsys Inc has been completely refurbished and modernized
  New RF tubes and coax; digital control systems; new vacuum system;
  ~0.9 MeV deuterons, < 2% duty factor, ~2 mA peak current, ~50 uA avg current
  Beam spot using permanent focusing quadru- pole magnets is ~1 cm at 2 m from RFQ exit
l
50us, 2mA beam pulse
Time [10 us/div]
Voltage [1V/div] (Calibration 1V/mA)

## Slide 15: A compact LaBr3:Ce scintillator coupled to an Si-APD provides high-resolution gamma spectroscopy
AGNOSTIC particle detection must be performed in an extremely hostile environment ● High neutron flux (~1013 m-2 s-1) ● High magnetic fields (< 0.1 T) ● Mechanical shock (~200 g) ● Compact geometry (~cm) ● High counts rates (<105 / s)

## Slide 16: A compact LaBr3:Ce scintillator coupled to an Si-APD provides high-resolution gamma spectroscopy
A photo of the LS detector next to a pencil for scale (left) and within its reentrant cartridge (right)
AGNOSTIC particle detection must be performed in an extremely hostile environment ● High neutron flux (~1013 m-2 s-1) ● High magnetic fields (< 0.1 T) ● Mechanical shock (~200 g) ● Compact geometry (~cm) ● High counts rates (<105 / s)
A 0.9 x 0.9 x 3.5 cm LaBr3:Ce crystal coupled to a Hamamatsu silicon avalanche photodiode in a ruggedized stainless steel housing was fabricated by Saint-Gobain for AIMS

## Slide 17: A compact LaBr3:Ce scintillator coupled to an Si-APD provides high-resolution gamma spectroscopy
A photo of the LS detector next to a pencil for scale (left) and within its reentrant cartridge (right)
AGNOSTIC particle detection must be performed in an extremely hostile environment ● High neutron flux (~1013 m-2 s-1) ● High magnetic fields (< 0.1 T) ● Mechanical shock (~200 g) ● Compact geometry (~cm) ● High counts rates (<105 / s)
A 0.9 x 0.9 x 3.5 cm LaBr3:Ce crystal coupled to a Hamamatsu silicon avalanche photodiode in a ruggedized stainless steel housing was fabricated by Saint-Gobain for AIMS
Energy resolution typically a factor of 2 to 3 better than NaI(Tl) detector; photo- peak statistics excellent despite 30x smaller sensitive volume than NaI(Tl) detector
LS Detector response to 661.7 keV gammas

## Slide 18: Advanced digital detector waveform acquisition and post-processing enable AIMS particle detection
Particle (n/γ)
discrimination
Spectroscopy
Pile-up
deconvolution
Pulse
inspection
Digital
memory
Waveform
Digitizers
With DPP
Baseline
restorers
Timers
Integrat-
ors
Discrimin-
ators
= Digital operations replace analog electronics
= Digital memory enables off-line analysis
LaBr3
EJ301

## Slide 19: Advanced digital waveform acquisition and post-processing enable AIMS particle detection
Pulse
inspection
Spectroscopy
Pile-up
deconvolution
Particle (n/γ)
discrimination
Gammas
Neutrons
Digital
memory
Waveform
Digitizers
With DPP
Baseline
restorers
Timers
Integrat-
ors
Discrimin-
ators
= Digital operations replace analog electronics
= Digital memory enables off-line analysis
LaBr3
EJ301

## Slide 20: Advanced digital detector waveform acquisition and post-processing enable AIMS particle detection
Pulse
inspection
Particle (n/γ)
discrimination
Pile-up
deconvolution
Spectroscopy
Digital
memory
Waveform
Digitizers
With DPP
Baseline
restorers
Timers
Integrat-
ors
Discrimin-
ators
= Digital operations replace analog electronics
= Digital memory enables off-line analysis
LaBr3
EJ301

## Slide 21: Mockup first-wall measurements were conducted with AIMS in the laboratory
PFCs
Neutrons
Gamma rays
C-Mod vacuum vessel
Gamma rays
~900 keV
Deuterons
LaBr3-SiAPD gamma detector
Experiments to simulate the actual PWI measurements on C-Mod were conducted ex-situ in the laboratory
Beam: ~0.9 MeV Deuterons 0.1% duty cycle 50 us bunches 30 Hz rep rate
Target: C-Mod Molybdenum PFC Tiles from inner wall
Objective: Validate ability to monitor wall conditions by quantifying boron and oxygen isotopes

## Slide 22: AIMS “proof-of-principle” measurements in the laboratory & C-Mod

## Slide 23: ACRONYM (built on GEANT4) provides full 3-D particle tracking  quantitative interpretation of gamma/neutron spectra through synthetic diagnostics of detectors

## Slide 24: Key issues to consider for interfacing AIMS
Port availability for both RFQ and the detectors.
Footprint of RF + HV cabinets that drive RFQ
Beam steering capacity (port + cw B field) and optimization
Radiation safety: RFQ operations require shielding  interlocks to cell access.
Interface to operations: RFQ operations + detectors
Lithium detection schemes to be explored by MIT team
  Know that both Li-6 and Li-7 have high Q nuclear reactions with D but we have to scope detection methods for gammas and neutrons.

## Slide 25: Questions