1.0   PURPOSE

This test procedure describes the method for testing the NSTX Thomson Scattering APD's high voltage power supply which is located on the Preamp/HV Supply printed circuit board.

2.0   SCOPE

The Responsible Manager will specify the Qualified Test Person who will perform the testing and a CPR qualified Safety Watch who will monitor the testing of the power supply.

Power will be applied to the input of the High Voltage Power Supply only after the printed circuit board is mounted in a grounded aluminum box and the high voltage monitoring meter is mounted inside of a Lucite enclosure.

3.0   PREREQUISITES

3.1  Qualified Test Person:

Signature  ___________________, Date:  ________

3.2  CPR qualified Safety Watch:

Signature  ___________________, Date:  ________

3.3  Installation of a "WARNING High Voltage May Be Present-Authorized Personnel Only" sign and a safety-barrier rope at entrance of lab test area is required.

3.4  Class O high voltage gloves shall be worn when trouble shooting an energized circuit.

4.0   REQUIRED EQUIPMENT

4.1  The following equipment is required to perform the high voltage performance test:

a) Digital Voltmeters (DVM), Fluke 8060A or equivalent - Qty = 2

b) Lab supplies to furnish +/- 5.0 Volts, +13.5 Volts and logic 5 Volts

c) Voltage Calibration Source, Datel DVC 8500 or equivalent

d) Small screw driver

5.0   PROCEDURE FOR TESTING PREAMP HIGH VOLTAGE POWER SUPPLY

5.1  Setup test circuit per figure 1.  Note: Preamp/HV Supply PC board is not install at this point.

Check step completed  ______

5.2  Select PC board to be tested.

a)  Record preamp/HV Supply PC board S/N.		       S/N  _______

b)  Check jumper setup: B - B1, V - V2 and T1 - T2.

c)  Adjust pot R45 fully CW.

d)  Adjust the Auto0/Hold test switch to the Auto0 position.

e)  Check that connections E13 to terminal "Vctr" of high voltage power supply module (HVPS) and E14 to HVPS terminal "+Vin" are made.

Check step completed  ______

5.3  Before installing PC board into test circuit, turn ON power supplies and check supply outputs: +5.0V +/- 0.1V, -5.0V +/- 0.1V, +13.5V +/- 0.3V and logic +5.0V +/- 0.3V.  Turn OFF power supplies.

Check step completed  ______

5.4  Initiate a Safety Watch and check that the "WARNING High Voltage May Be Present-Authorized Personnel Only" sign and a safety-barrier rope at entrance of lab test area are installed.

Check step completed  ______

5.5  Install PC board to be tested in grounded aluminum box.  Connect 15 contact "D" test connector to the PC board.

Check step completed  ______

5.6  Turn on the high voltage monitoring DVM and set its range for 1000V DC. Install the high voltage monitoring DVM inside of a plastic Lucite enclosure.  Connect high voltage monitoring DVM leads to E18 (HI) and E15 (RTN) on the PC board.

Check step completed  ______

5.7  Install grounded cover on aluminum box.

Check step completed  ______

5.8 Adjust Voltage Calibration Source digi-switches for 0.000V, (Control Voltage).

Check step completed  ______

5.9  Turn ON the +/- 5V, 13.5V and logic +5V power supplies, plus the Voltage Calibration Source.

Check step completed  ______

5.10  Record the HVPS output (E18 - E15) and HV telemetry output for Control Voltage = 0.000V.

HVPS output               _________V,   ( < 4V),

HV telemetry output   _________V,   (≈0.04V)

Control Voltage          _________V,   (0.000V)

5.11  Adjust  Control Voltage for HVPS output = 300.0V +/- 0.3V.  Record the HVPS output, the HV telemetry output and the Control Voltage.

HVPS output               _________V,   (300.0V +/- 0.3V),

HV telemetry output   _________V,   (≈3.000V)

Control Voltage          _________V,   (≈2.43V)

5.12  Adjust  Control Voltage for HVPS output = 500.0V +/- 0.5V.  Record the HVPS output, the HV telemetry output and the Control Voltage.

HVPS output               _________V,   (500.0V +/- 0.5V),

HV telemetry output   _________V,   (≈5.000V)

Control Voltage          _________V,   (≈4.15V)

5.13  Adjust Control Voltage for 0.000V.  Wait until the HVPS output < 10V and HV telemetry output < 0.1V, then turn OFF the +/- 5V, 13.5V and logic +5V power supplies, plus the Voltage Calibration Source.  Before preceding to the next step, observe that HVPS output < 1V and HV telemetry output < 0.01V. 										Check step completed  ______

5.14  Remove the cover on aluminum box.  Remove the high voltage monitoring DVM leads from the PC board being tested.  Reinstall the grounded cover on the aluminum box.

Check step completed  ______

5.15  Adjust Voltage Calibration Source digi-switches for 0.000V, (Control Voltage).  Turn ON the +/- 5V, 13.5V and logic +5V  power supplies, plus the Voltage Calibration Source.

Check step completed  ______

5.16  Adjust  the Control Voltage for the value recorded in step 5.11. Record the HV telemetry output and the Control Voltage.

HV telemetry output   _________V,   (≈3.000V)

Control Voltage          _________V,   (≈2.43V)

5.17  Adjust  the Control Voltage for the value recorded in step 5.12. Record the HV telemetry output and the Control Voltage.

HV telemetry output   _________V,   (≈5.000V)

Control Voltage          _________V,   (≈4.15V)

5.18  Adjust  Control Voltage for 0.000V.  Wait until the HV telemetry output < 0.1V, then turn OFF the +/- 5V, 13.5V and logic +5V power supplies, plus the Voltage Calibration Source.  Before preceding to the next step, observe that HV telemetry output < 0.01V.

Check step completed  ______

5.19  Remove the cover on aluminum box and remove the tested PC board.

Check step completed  ______

5.20  Remove the "WARNING High Voltage May Be Present-Authorized Personnel Only" sign and a safety-barrier rope at entrance of lab test area if testing of PC boards are completed.

Check step completed  ______