National Spherical Tokamak Experiment

NSTX

SYSTEM REQUIREMENTS DOCUMENT

DIAGNOSTICS

NSTX-SRD-4X-016

Revision 0

May 2, 1996

________________________________

Robert Kaita

WBS DIAGNOSTICS Manager

________________________________

Charles Neumeyer

NSTX Project Engineering Manager

________________________________

Masa Ono

NSTX Project Manager

| Revision | Date | TRB | Description of Change |
| --- | --- | --- | --- |

Introduction	4

DIAGNOSTIC System Requirements	6

1.0 Functions	6

1.1 Day One Diagnostics	6

1.2 Baseline Diagnostics	6

1.3 Upgrade Diagnostics	7

2.0 Configuration Requirements	8

2.1 Diagnostic Port Requirements	8

3.0 Operational Requirements	9

3.1 Data Acquisition System Requirements	9

List of Tables

D-1	day one and baseline nstx diagnostics summary	5

D-1.2-1	baseline nstx single channel diagnostics	7

D-1.2-2	baseline nstx profile diagnostics	7

D-1.3-1	upgrade nstx diagnostics	8

D-2.1-1	nstx diagnostic port requirements	9

APPENDIX

A.1 Special NSTX Diagnostic Needs	11

A.1.1 Divertor Diagnostics	11

A.1.2 Current Profile Measurements	11

A.1.3 Electron Temperature Profile Measurements	12

Introduction

The purpose of plasma diagnostics is to provide information on discharge parameters, d guide NSTX operations for optimized performance. Those diagnostics which are essential for machine operations are called Day One diagnostics, and they are indicated with an asterisk in Table D-1.  The abbreviations used in the table are as follows.

1) TVTS: Multichannel Thomson Scattering System (electron temperatures and densities)

2) CHERS: Charge-Exchange Recombination Spectroscopy (ion temperatures and rotation)

3) MSE: Motional Stark Effect Polarimetry (current profiles)

Although Mirnov coils are not critical for machine operations, they are also listed as Day One diagnostics because they have to be installed as part of the NSTX assembly process.

After plasma operations on NSTX begin, the near-term emphasis will be on detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. Equipment which is presently available at PPPL will be used, and this is listed also in Table D-1. Since current profile measurements are particularly important, special emphasis will be placed on installing the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam.

The long-term objective is to upgrade and expand the diagnostic set for the study  of fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics will provide input for advanced plasma control systems. Possible upgrade diagnostics that currently exist at PPPL include a “DensPak” array of Mirnov coils, a two-dimensional soft X-ray imaging system, beam emission spectroscopy, amplitude modulation reflectometry (in collaboration with the Oak Ridge National Laboratory and the EURATOM/CIEMAT group in Madrid, Spain), a multichannel fast neutral analyzer, and fast and slow neutron counters. New concepts and systems developed by the national NSTX team to implement these goals also will be encouraged.

Table D-1: Day One and Baseline Diagnostics Summary

| Multichannel TVTS (single pulse) | Te(r), ne(r) |
| --- | --- |

| CHERS (with neutral probe beam) | Ti(r), vφ(r) |
| --- | --- |

| MSE (with neutral probe beam) | j(r) |
| --- | --- |

| Multichannel bolometer | Prad(r) |
| --- | --- |

| Visible continuum array | Zeff(r) |
| --- | --- |

| 1 mm microwave interferometer | nel |
| --- | --- |

| Survey spectrometer (SPRED) | Impurities |
| --- | --- |

| Soft X-ray imaging system | Fluctuations |
| --- | --- |

| Mirnov coils* | Fluctuations |
| --- | --- |

| Flux loops | Fluctuations |
| --- | --- |

| Rogowski coils* | Ip |
| --- | --- |

| Flux Loops* | Magnetics |
| --- | --- |

| Hα detectors | Edge recycling |
| --- | --- |

| Langmuir probes | Edge parameters |
| --- | --- |

| Visible camera* | Plasma control |
| --- | --- |

| IR camera* | Heat loads |
| --- | --- |

| Residual gas analyzers | In-vessel gases |
| --- | --- |

WBS 4X -  Diagnostic System Requirements

1.0 Functions

1.1 Day One Diagnostics

The diagnostics which are essential for machine operations are called Day One diagnostics, and they are indicated with an asterisk in Table D-1. Although Mirnov coils are not critical for machine operations, they are also listed as Day One diagnostics because they have to be installed as part of the NSTX assembly process.

The magnetic diagnostics shall be similar to the coil/loop sets installed on the Princeton Beta Experiment-Modification. The components shall be as follows.

1) Plasma current Rogowski coils: 2 toroidal locations.

2) Eddy current Rogowski coils: 27 on passive stabilizers.

3) Voltage loops: 9 poloidal locations.

4) Bθ, Bρ coils: 20 locations.

5) Mirnov coils: 120 locations.

6) Flux loops: 20 locations.

7) Segmented loops: 10 locations.

All coils/loops shall be installed in a way that minimizes plasma damage. This is a particular concern for coils/loops to be mounted on the center column. All coils/loops shall also be compatible with bakeout requirements as specified in WBS 33.

The diamagnetic loop shall be an upgrade diagnostic. Aligning it is time consuming, and the TF coils can be used instead in the high-β NSTX plasmas.

Visible and infrared (IR) cameras shall be installed with adequate views of the main plasma chamber (including the passive stabilizer plates) and the divertor region.

Thermocouples shall be placed on the backing plates for the plasma facing tiles in the main plasma chamber (including the passive stabilizer plates) and the divertor region.

Additional divertor and scrapeoff layer diagnostics shall be installed as upgrade diagnostics. They are described in Appendix A.1.1.

1.2 Baseline Diagnostics

The near-term emphasis will be on detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. The magnetics and machine operations diagnostics (including cameras and thermocouples) are Day One diagnostics that are described in Section 1.1. The line-averaged density, impurity, and edge diagnostics in the baseline NSTX diagnostic set are listed in Table D-1.2-1. The baseline profile diagnostics for NSTX are given in Table D-1.2-2.

Table D-1.2-1: Baseline NSTX Single Channel Diagnostics

| Category | System | Status |
| --- | --- | --- |

| Electron Density | 1 mm Microwave Interferometer | some components in storage; may need to purchase replacement parts |
| --- | --- | --- |

| Impurities and Edge Diagnostics | Bolometers Hα Detectors | components in storage |
| --- | --- | --- |

|  | Vacuum Ultraviolet Spectrometer (“SPRED”) | obtain from TFTR |
| --- | --- | --- |

|  | Langmuir Probes | (readily available) |
| --- | --- | --- |

|  | Gas Injection System | (readily available) |
| --- | --- | --- |

|  | Residual Gas Analyzer | (readily available) |
| --- | --- | --- |

Table D-1.2-2: Baseline NSTX Profile Diagnostics

| System | Status | No. Pts. | Δr (cm)) | Δt (ms) |
| --- | --- | --- | --- | --- |

| TVTS | on TFTR | 55 | 1.2 | (single) |
| --- | --- | --- | --- | --- |

| CHERS | spectrometer from TFTR; neutral probe beam on PBX-M | 12 | 3 | 5 |
| --- | --- | --- | --- | --- |

| MSE | detectors on TFTR; neutral probe beam on PBX-M | 10 | 3 | 5 |
| --- | --- | --- | --- | --- |

| Tangential Bolometer | components in storage | 15 | 3 | 3 |
| --- | --- | --- | --- | --- |

| Visible Continuum | components in storage | 15 | 4 | 3 |
| --- | --- | --- | --- | --- |

| Soft X-Ray Imaging | components in storage | 32 | 2 - 3 | .003 |
| --- | --- | --- | --- | --- |

1.3 Upgrade Diagnostics

The long-term objective is to upgrade and expand the diagnostic set for the study  of fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics will provide input for advanced plasma control systems. Possible upgrade diagnostics that currently exist at PPPL are given in Table D-1.3-1. Beam emission spectroscopy requires the diagnostic neutral probe beam or a neutral heating beam. The fast ion and neutron diagnostics are easily movable, but are not needed prior to the installation of the neutral heating beams.

Table D-1.3-1: Upgrade NSTX Diagnostics

| Measurement/System | Status |
| --- | --- |

| Magnetic Fluctuations/ “DensPak” Mirnov Coils | on PBX-M |
| --- | --- |

| Diamagnetic Loop | (new installation) |
| --- | --- |

| Tomography/Soft X-Ray Imaging System | electronics on PBX-M; phosphor screen in storage |
| --- | --- |

| Fluctuations/AM Reflectometry | antennas and waveguide on PBX-M; electronics at General Atomics (DIII-D) |
| --- | --- |

| Fluctuations/Beam Emission Spectroscopy | on PBX-M and TFTR |
| --- | --- |

| Fluctuations/Multichannel Radiometer | used on PBX-M and in storage at MIT |
| --- | --- |

| Fluctuations/Fast Scanning Edge Probe | on PBX-M but easily movable between machine areas |
| --- | --- |

| Fast Ions/Multichannel Charge-Exchange Analyzer | in storage |
| --- | --- |

| Neutrons/Gas Neutron Counters | on PBX-M |
| --- | --- |

| Neutrons/Plastic Scintillator Detectors | on PBX-M |
| --- | --- |

In addition, the national NSTX team will be encouraged to develop new diagnostic concepts and systems to implement the program objectives of detailed physics understanding of high-performance plasmas and advanced plasma control.

2.0 Configuration Requirements

2.1 Diagnostic Port Requirements

The plasma diagnostics on NSTX shall connect to the vacuum vessel and support structure (WBS 12) at the machine interface vacuum valve and/or the points of mechanical support for the diagnostic inside the vacuum vessel. The assumptions concerning the vacuum vessel ports on NSTX are as follows.

1)	Eight midplane ports available: four ports needed for neutral beam

injectors and dumps.

2)	Four out of eight planned RF antennas will cross midplane in space

between ports.

3)	Remaining four RF antennas will still permit use of about half of

midplane ports they cross.

4)	Straps for RF antennas will use off-midplane ports.

With these constraints, the port requirements are given in Table D-2.1-1.

Table D-2.1-1: NSTX Diagnostic Port Requirements

| Location | Diagnostic System |
| --- | --- |

| Midplane port and dome port | Multichannel TVTS (single pulse) Hα detectors |
| --- | --- |

| Midplane port | CHERS MSE |
| --- | --- |

| Midplane port | Dump for TVTS laser and neutral probe beam |
| --- | --- |

| Midplane port | Multichannel bolometer Visible continuum array |
| --- | --- |

| Midplane port | Diagnostic neutral probe beam Survey spectrometer (SPRED) |
| --- | --- |

| Midplane port and dome port | Soft X-ray imaging system 1 mm microwave interferometer |
| --- | --- |

| Midplane port | Hα detectors Langmuir probes Residual gas analyzer |
| --- | --- |

| Midplane port | Visible camera IR camera |
| --- | --- |

All connections between the diagnostics and the vacuum vessel shall be compatible with the conditions during vacuum vessel bakeout (WBS 33).

The vacuum system, if required, for each diagnostic shall be provided by the individual or group responsible for fabricating it. This includes, for example, forepumps, turbopumps, cryopumps, and associated controllers. The exhaust outlets for forepumps shall be provided as part of the NSTX vacuum pumping system (WBS 31).

3.0 Performance & Operational Requirements

3.1 Data Acquisition System Requirements

Planning and costing for each diagnostic shall include the device control and data acquisition modules that interface with the data highway of the NSTX Data Acquisition System (WBS 62).

The tabulated data acquisition signals require sampling rates ranging from 100 Hz to 1 MHz and shall be used to meet NSTX diagnostics data acquisition requirements.

To establish requirements for Central I&C, each Data Acquisition signal is included in one of the following general categories:

Low Speed ==> Data Acquisition signals requiring sample rates >100 Hz and ≤2 KHz.

Moderate Speed ==> Data Acquisition signals  requiring sample rates >2 KHz and ≤100 KHz.

High Speed ==> Data Acquisition signals  requiring sample rates >100 KHz and ≤1 MHz.

The Average Sampling Time for each signal category (per NSTX pulse) is also included in the table below.

Table D-3.1-1: Data Acquisition Requirements

|  | Number of Signals & Average Sampling Time per Pulse |
| --- | --- |

| NSTX Diagnostics | Low Speed >100Hz & ≤2KHz | AVG SMPL Time (sec/ pulse) | Mod. Speed >2KHz & ≤100 KHz | AVG SMPL Time (sec/ pulse) | High Speed >100 KHz & ≤1MHz | AVG SMPL Time (sec/ pulse) |
| --- | --- | --- | --- | --- | --- | --- |

| Day One Diagnostics | 250 | 1 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |

| Baseline Diagnostics |  |  | 150 | 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |

| Upgrade Diagnostics |  |  | 50 | 5 | 100 | 5 |
| --- | --- | --- | --- | --- | --- | --- |

| TOTALS | 250 |  | 200 |  | 100 |  |
| --- | --- | --- | --- | --- | --- | --- |

The following will be assumed to minimize costs.

1) The existing PPPL computer network shall be used. The present bus architecture connects alpha workstations, X-terminals, etc. in a network that includes collaboration links. The network supports distributed diagnostic control and data acquisition and centralized data management.

2) The interface between the data highway and the diagnostic device control and data acquisition modules shall be a hybrid that can accommodate existing CAMAC equipment and hardware based on newer standards.

The Day One and baseline diagnostics are expected to generate approximately 3.2 megabytes (MB) of data for each one-second pulse. The upgrade diagnostics are expected to generate approximately 33.2 MB for each one-second pulse. These estimates are comparable to the values for PBX-M at the end of its most recent operating period. They should be increased by about a factor of five when the NSTX pulse duration is lengthened to five seconds.

A.1 Special NSTX Diagnostic Needs

A.1.1  Divertor Diagnostics

The following are the edge, scrapeoff layer, and divertor diagnostics that are possible on NSTX.

1) Bolometry (baseline diagnostic): tangential array in horizontal midplane and divertor views.

2) Plasma TV Camera (Day One diagnostic): tangential views of main plasma chamber and divertor regions.

3) IR Camera (Day One diagnostic): tangential views of main plasma chamber and divertor regions.

The Langmuir probes and Hα detectors are straightforward diagnostics. Pressure gauges are not included in this list, since they are considered part of the Auxiliary Systems (Vacuum Pumping System) WBS Element (WBS 31). It is critical, however, that there are enough of them to monitor the pressure in the divertor as well as the main plasma chamber.

In addition to the thermocouple arrays listed as Day One diagnostics on the backing plate structure for the plasma facing tiles, divertor probe arrays and Rogowski coils on backing plate supports shall also be installed. These diagnostics shall either be available at the start of plasma operations if they are part of the Tokamak Systems (Plasma Facing Components) WBS element (WBS  11), or installed later as upgrade diagnostics.

A.1.2  Current Profile Measurements

Because of the importance of current profile measurements, installation of the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam has high priority. The special issues that must be addressed are as follows.

•	Diagnostic Neutral Probe Beam (NPB):

1) Operational reliability - The NPB was commissioned at the Culham Laboratory (UK) for operation with 100 ms pulses at 80 kV. This was never achieved at PPPL because of power supply difficulties. They should be corrected as part of the task of moving the beamline to the NSTX platform (test cell) and installing a new control console in the NSTX control room.

2) Long pulse operation - The existing power supplies (when working reliably) permit 1 s operation. New grids are needed to accommodate this pulse length. Longer (5 s) NSTX discharges can then be accommodated by operating the NPB in pulsed mode (e. g., 20 ms pulses every 100 ms for 50 time points in 5 s) but major source and power supply upgrades are needed for 5 s DC operation.

3) Availability of the neutral probe beam - PBX-M will require the diagnostic neutral probe beam (NPB) for at least a year of operations, while a new scheme to instrument one of the heating beams for motional Stark effect (MSE) polarimetry is being tested. The use of a heating beam on PBX-M is speculative because it has a lower injection energy than the NPB, so we need to determine if the improved signal from the higher heating beam current will compensate for it. If PBX-M resumes operation in FY99, the schedule shall be to move the NPB to NSTX at the beginning of FY2000 at the earliest, and begin current profile measurements in the middle of that fiscal year. This will avoid having to develop MSE diagnostics simultaneously on NSTX (see below) and PBX-M, which are both really new because of the novel physics constraints.

•	Motional Stark Effect (MSE) Diagnostic:

1) Feasibility of measurement - The magnitude of the NSTX toroidal field relative to the NPB injection energy and the NPB trajectory relative to the MSE sightlines must be evaluated to determine if the Stark-shifted line can be resolved. Low field, low aspect ratio, and port access are all potential difficulties.

A.1.3  Electron Temperature Profile Measurements

Electron temperature and density profiles will be critical to the understanding of the confinement and stability of NSTX plasmas. The initial phase of NSTX operations will be characterized by short pulses and high repetition rates. This is similar to the present situation on PBX-M, where the temporal evolution of electron temperature and density profiles is determined from multiple discharges using a TV Thomson scattering (TVTS) system with a single-pulse laser. In the long term, however, it is desirable to measure the electron temperature and density profiles as a function of time in a single discharge. This is particularly true when the NSTX pulse lengths are extended to 5 s, and profile information is also needed for feedback applications. Two options for a multi-time-point TVTS system are summarized as follows.

1) Install multipulse glass laser and upgrade imaging and computer systems at a cost of ≈ $0.5M.

2) Install completely new Nd:YAG laser-based system at a cost of ≈ $1.5M

The first option is attractive because of its lower cost, but the overall diagnostic concept has not been tested. The cost for the second possibility is much higher, but the system has been demonstrated for profile measurements in multi-second discharges on the DIII-D tokamak. As a system in the upgrade set, the choice of which TVTS option to pursue will depend on when the NSTX program needs the diagnostic and the funds available for it.