National Spherical Tokamak Experiment

NSTX

SYSTEM REQUIREMENTS DOCUMENT

DIAGNOSTICS

NSTX-SRD-4X-016

Revision 0

July 16, 1996

________________________________

Robert Kaita

WBS DIAGNOSTICS Manager

________________________________

Charles Neumeyer

NSTX Project Engineering Manager

________________________________

Masayuki Ono

NSTX Project Manager

| Revision | Date | TRB | Description of Change |
| --- | --- | --- | --- |

Introduction	4

DIAGNOSTIC System Requirements	6

1.0 Functions	6

1.1 Day One Diagnostics	6

1.2 Baseline Diagnostics	6

1.3 Upgrade Diagnostics	6

2.0 Configuration Requirements and Essential Features	7

2.1 Diagnostic Port Requirements	7

2.2 Requirements for Internal Diagnostics	7

2.3 Diagnostic Vacuum System Requirements	8

2.4 Electrical Power and Isolation Requirements	8

List of Tables

D-1	day one and baseline nstx diagnostics summary	5

D-1.3-1	upgrade nstx diagnostics	6

D-2.1-1	nstx diagnostic port requirements	7

APPENDIX

A.1 Special NSTX Diagnostic Needs	9

A.1.1 Current Profile Measurements	9

A.1.2 Electron Temperature Profile Measurements	10

Introduction

The purpose of plasma diagnostics is to provide information on discharge parameters to characterize NSTX plasmas and guide NSTX operations for optimized performance. Those diagnostics which are essential for machine operations are called Day One diagnostics, and they are indicated with an asterisk in Table D-1.

After plasma operations on NSTX begin, the near-term program emphasis shall be on detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. Equipment which is presently available at PPPL shall be used, and this is listed also in Table D-1. Since current profile measurements are particularly important, special emphasis shall be placed on installing the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam.

The long-term objective is to upgrade and expand the diagnostic set for the study  of fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics shall provide input for advanced plasma control systems. Possible upgrade diagnostics that currently exist at PPPL include a “DensPak” array of Mirnov coils, a two-dimensional soft X-ray imaging system, beam emission spectroscopy, amplitude modulation reflectometry (in collaboration with the Oak Ridge National Laboratory and the EURATOM/CIEMAT group in Madrid, Spain), a multichannel fast neutral analyzer, and fast and slow neutron counters. New concepts and systems developed by the national NSTX team to implement these goals also shall be encouraged.

Table D-1: Day One* and Baseline Diagnostics Summary

| System | Function |
| --- | --- |

Plasma current Rogowski coils*

Total plasma current

Eddy current Rogowski coils*

Vertical position control

| Voltage loops* | Poloidal flux for plasma control |
| --- | --- |

Bq, Br coils*

Magnetic fluctuations

| Mirnov coils* | Plasma control/magnetic fluctuations |
| --- | --- |

| Flux loops* | Toroidal flux for plasma control |
| --- | --- |

| Segmented loops* | Flux measurements on passive plates |
| --- | --- |

| Visible TV camera* | External shape for plasma control |
| --- | --- |

| IR camera* | Heat loads |
| --- | --- |

| Residual gas analyzers* | In-vessel gases |
| --- | --- |

| Multichannel TVTS1 (single pulse) | Electron temperature and density |
| --- | --- |

| CHERS2 (with neutral probe beam) | Ion temperature and toroidal rotation |
| --- | --- |

| MSE3 (with neutral probe beam) | Plasma current density profile |
| --- | --- |

| Multichannel bolometer | Radiated power profile |
| --- | --- |

| Visible continuum array | Zeff(r) |
| --- | --- |

| 1 mm microwave interferometer | Line-integrated plasma density |
| --- | --- |

| Survey spectrometer (SPRED) | Plasma impurities |
| --- | --- |

| Hα detectors | Edge recycling |
| --- | --- |

| Langmuir probes | Edge parameters |
| --- | --- |

Soft X-ray imaging system

Plasma instabilities and fluctuations

The abbreviations used in the table are as follows.

1TVTS: Multichannel Thomson Scattering System

2CHERS: Charge-Exchange Recombination Spectroscopy

3MSE: Motional Stark Effect Polarimetry

WBS 4X -  Diagnostic System Requirements

1.0 Functions

1.1 Day One Diagnostics

The function of Day One diagnostics shall be to support basic machine operations. They are listed in Table D-1.

1.2 Baseline Diagnostics

The function of the baseline diagnostics shall be to support the near-term emphasis of the NSTX program. This shall be effected by detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. The baseline diagnostics are listed in Table D-1.

1.3 Upgrade Diagnostics

The function of the upgrade diagnostics shall be to support the long-term NSTX program objective of studying fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics shall provide input for advanced plasma control systems. Possible upgrade diagnostics are given in Table D-1.3-1.

Table D-1.3-1: Upgrade NSTX Diagnostics

System

Function

“DensPak” Mirnov Coils

Mode identification/magnetic fluctuations

| Diamagnetic Loop | Stored energy |
| --- | --- |

Soft X-Ray Imaging System

Tomographic reconstruction of plasmas

AM Reflectometry

Density profiles/fluctuations

Beam Emission Spectroscopy

Density fluctuations

| Multichannel Radiometer | Fluctuations/“ballooning modes” |
| --- | --- |

| Fast Scanning Edge Probe | Fluctuations/edge plasma profiles |
| --- | --- |

Multichannel Charge-Exchange Analyzer

Fast ion distributions

Gas Neutron Counters

Thermal neutrons

Plastic Scintillator Detectors

Fast neutron fluctuations

a. Divertor probe arrays and Rogowski coils on backing plate supports shall be installed as upgrade diagnostics.

b. The national NSTX team shall be encouraged to develop new diagnostic concepts and systems to implement the program objectives of detailed physics understanding of high-performance plasmas and advanced plasma control.

2.0 Configuration Requirements and Essential Features

2.1 Diagnostic Port Requirements

a. The plasma diagnostics on NSTX shall connect to the vacuum vessel and support structure (WBS 12) at the machine interface vacuum valve and/or the points of mechanical support for the diagnostic inside the vacuum vessel. Four large midplane ports and at least three dome ports shall be available for diagnostics. The port utilization shall be as given in Table D-2.1-1.

Table D-2.1-1: NSTX Diagnostic Port Requirements

| Location | Diagnostic System |
| --- | --- |

Large midplane port

Dome port above midplane for XIS

Dome port below midplane for TVTS

Multichannel TVTS (single pulse)

Soft X-ray imaging system (XIS)

Survey spectrometer (SPRED)

Large midplane port

Multichannel bolometer

Single channel bolometers

Visible continuum array

| Large midplane port | Visible TV camera IR camera Hα detectors |
| --- | --- |

| Large midplane port | CHERS MSE 1 mm microwave interferometer Langmuir probes |
| --- | --- |

Large midplane port (for neutral heating beam during NSTX upgrade phase)

Diagnostic neutral probe beam (will be removed during NSTX upgrade phase)

| Dome port | Residual gas analyzer |
| --- | --- |

b. The neutral probe beam shall use a large midplane port that is reserved for a neutral heating beam during the NSTX upgrade phase. The neutral heating beam shall perform the function of the neutral probe beam during the NSTX upgrade phase.

c. The visible TV and infrared (IR) cameras shall have tangential views of main plasma chamber and divertor regions

d. The multichannel bolometer shall consist of a tangential array in horizontal midplane.

e. The two single channel bolometers shall have views of the upper and lower divertors.

f. The dumps for the TVTS laser and the neutral probe beam shall be located between ports.

g. All connections between the diagnostics and the vacuum vessel shall be compatible with a 150o C vacuum vessel bakeout temperature.

2.2 Requirements for Internal Diagnostics

a. All coils, loops, and other diagnostics mounted inside the vacuum vessel shall be compatible with a 150o C vacuum vessel bakeout temperature.

b. All coils, loops, and other diagnostics mounted inside the vacuum vessel shall be installed in a way that minimizes plasma damage. This includes protection against high voltages and large heat loads, particularly for components located on the center column.

c. The number of coils and loops shall be as follows.

Plasma current Rogowski coils: up to 2 toroidal locations.

Eddy current Rogowski coils: up to 27 on passive stabilizers.

Voltage loops: up to 9 poloidal locations.

Bθ, Bρ coils: up to 20 locations.

Mirnov coils: up to 120 locations.

Flux loops: up to 20 locations.

Segmented loops: up to 10 locations.

d. Thermocouples shall be placed on the backing plates of the plasma facing tiles in the main plasma chamber (including the passive stabilizer plates) and the divertor region.

2.3 Diagnostic Vacuum System Requirements

a. The vacuum system, if required, for each diagnostic shall be included in the diagnostic fabrication task. This includes, for example, forepumps, turbopumps, cryopumps, and associated controllers.

b. The exhaust outlets for forepumps shall be provided as part of the NSTX vacuum pumping system (WBS 31).

2.4 Electrical Power and Isolation Requirements

a. All diagnostics shall have electrical isolation from the vacuum vessel that is rated for 5 kV.

b. All connections to and from diagnostic equipment in the NSTX test cell shall be with optical fiber. All connections between the data acquisition systems for each diagnostic shall be with optical fiber.

c. All diagnostic signals shall be isolated via optical or magnetic means prior to exiting the test cell boundary. The isolation shall be rated to withstand a one minute AC hipot test at 20 kV AC rms.

d. There shall be two independent AC power distribution systems. One 60 Hz, 120 V up to 480 V system shall be designated for diagnostic signal processing, data acquisition, and control subsystems. The other 60 Hz, 120 V up to 480 V system shall be used to power motors and other “noisy” or high power equipment.

e. Each diagnostic shall be independently grounded to a single point ground as determined by the NSTX machine grounding specifications.

A.1 Special NSTX Diagnostic Needs

A.1.1  Current Profile Measurements

Because of the importance of current profile measurements, installation of the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam has high priority. The special issues that must be addressed are as follows.

•	Diagnostic Neutral Probe Beam (NPB):

1) Operational reliability - The NPB was commissioned at the Culham Laboratory (UK) for operation with 100 ms pulses at 80 kV. This was never achieved at PPPL because of power supply difficulties. They should be corrected as part of the task of moving the beamline to the NSTX platform (test cell) and installing a new control console in the NSTX control room.

2) Long pulse operation - The existing power supplies (when working reliably) permit 1 s operation. New grids are needed to accommodate this pulse length. Longer (5 s) NSTX discharges can then be accommodated by operating the NPB in pulsed mode (e. g., 20 ms pulses every 100 ms for 50 time points in 5 s) but major source and power supply upgrades are needed for 5 s DC operation.

3) Availability of the neutral probe beam - PBX-M will require the diagnostic neutral probe beam (NPB) for at least a year of operations, while a new scheme to instrument one of the heating beams for motional Stark effect (MSE) polarimetry is being tested. The use of a heating beam on PBX-M is speculative because it has a lower injection energy than the NPB, so we need to determine if the improved signal from the higher heating beam current will compensate for it. If PBX-M resumes operation in FY99, the schedule shall be to move the NPB to NSTX at the beginning of FY2000 at the earliest, and begin current profile measurements in the middle of that fiscal year. This will avoid having to develop MSE diagnostics simultaneously on NSTX (see below) and PBX-M, which are both really new because of the novel physics constraints.

•	Motional Stark Effect (MSE) Diagnostic:

1) Feasibility of measurement - The magnitude of the NSTX toroidal field relative to the NPB injection energy and the NPB trajectory relative to the MSE sightlines must be evaluated to determine if the Stark-shifted line can be resolved. Low field, low aspect ratio, and port access are all potential difficulties.

A.1.2  Electron Temperature Profile Measurements

Electron temperature and density profiles will be critical to the understanding of the confinement and stability of NSTX plasmas. The initial phase of NSTX operations will be characterized by short pulses and high repetition rates. This is similar to the present situation on PBX-M, where the temporal evolution of electron temperature and density profiles is determined from multiple discharges using a TV Thomson scattering (TVTS) system with a single-pulse laser. In the long term, however, it is desirable to measure the electron temperature and density profiles as a function of time in a single discharge. This is particularly true when the NSTX pulse lengths are extended to 5 s, and profile information is also needed for feedback applications. Two options for a multi-time-point TVTS system are summarized as follows.

1) Install multipulse glass laser and upgrade imaging and computer systems at a cost of ≈ $0.5M.

2) Install completely new Nd:YAG laser-based system at a cost of ≈ $1.5M

The first option is attractive because of its lower cost, but the overall diagnostic concept has not been tested. The cost for the second possibility is much higher, but the system has been demonstrated for profile measurements in multi-second discharges on the DIII-D tokamak. As a system in the upgrade set, the choice of which TVTS option to pursue will depend on when the NSTX program needs the diagnostic and the funds available for it.