National Spherical Tokamak Experiment

NSTX

SYSTEM DESIGN DESCRIPTION

DIAGNOSTICS

NSTX-SDD-4X-031

Revision 0

JULY 16, 1996

________________________________

Robert Kaita

WBS DIAGNOSTICS Manager

________________________________

Charles Neumeyer

NSTX Project Engineering Manager

________________________________

Masa Ono

NSTX Project Manager

| Revision | Date | TRB | Description of Change |
| --- | --- | --- | --- |

Introduction	1

DIAGNOSTICS design description	3

1.0 Design Description	3

1.1 System Design Features	3

1.1.1 Functions	3

1.1.1.1 Day One Diagnostics	3

1.1.1.2 Baseline Diagnostics	3

1.1.1.3 Upgrade Diagnostics	3

1.1.2 Configuration Description and Essential Features	4

1.1.2.1 Internal Diagnostics	4

1.1.2.2 Diagnostic Vacuum System	4

1.1.2.3 Electrical Power and Isolation	4

1.2 NSTX Diagnostics Layout	4

1.3 Component Design Features and System Performance	6

1.3.1 General Magnetic Diagnostics	6

1.3.1.1 Design Features	6

1.3.1.2 System Performance	6

1.3.2 Plasma TV Camera	9

1.3.2.1 Design Features	9

1.3.2.2 System Performance	9

1.3.3 IR Camera	9

1.3.3.1 Design Features	9

1.3.3.2 System Performance	9

1.3.4 Thomson Scattering (TVTS)	10

1.3.4.1 Design Features	10

1.3.4.2 System Performance	10

1.3.5 Charge-Exchange Recombination Spectroscopy (CHERS)	10

1.3.5.1 Design Features	10

1.3.5.2 System Performance	10

1.3.6 Motional Stark Effect Polarimetry (MSE)	11

1.3.6.1 Design Features	11

1.3.6.2 System Performance	11

1.3.7 Bolometry	11

1.3.7.1 Design Features	11

1.3.7.2 System Performance	12

1.3.8 Visible Bremsstrahlung Array 	12

1.3.8.1 Design Features	12

1.3.8.2 System Performance	12

1.3.9 Microwave Interferometer	12

1.3.9.1 Design Features	12

1.3.9.2 System Performance	13

1.3.10 Vacuum Ultraviolet Spectrometer (SPRED	13

1.3.10.1 Design Features	13

1.3.10.2 System Performance	13

1.3.11 Hα Detectors	13

1.3.11.1 Design Features	13

1.3.11.2 System Performance	14

1.3.12 Langmuir Probes	14

1.3.12.1 Design Features	14

1.3.12.2 System Performance	14

1.3.13 Soft X-Ray Diode Imaging System (XIS)	14

1.3.13.1 Design Features	14

1.3.13.2 System Performance	15

2.0 System Performance	15

3.0 Interfaces	15

3.1 Diagnostic Port Utilization	15

3.2 Data Acquisition System	16

4.0 Design Basis	17

4.1 Standards & Codes	17

4.2 Supporting Analysis	17

5.0 Design Implementation	17

5.1 Prototypes	17

5.2 Mock-ups	17

5.3 Component Procurement and Fabrication	17

5.4 System Assembly, Installation, and Testing	17

List of Tables

D-1	day one and baseline nstx diagnostics summary	2

D-1.1.1.3-1	upgrade nstx diagnostics	3

D-3.1-1	nstx diagnostic port summary	16

D-3.2-1	data acquisition summary	17

List of FIGURES

D-1.2-1	NSTX diagnostics layout	5

D-1.3.1-1	location of magnetic diagnostics	7

D-1.3.1-2	mounting concept for magnetic diagnostics	8

APPENDIX

A.1 Special NSTX Diagnostic Needs	18

A.1.1 Current Profile Measurements	18

A.1.2 Electron Temperature Profile Measurements	19

Introduction

The purpose of plasma diagnostics is to provide information on discharge parameters to characterize NSTX plasmas and guide NSTX operations for optimized performance. Those diagnostics which are essential for machine operations are called Day One diagnostics, and they are indicated with an asterisk in Table D-1.

After plasma operations on NSTX begin, the near-term program emphasis will be on detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. Equipment which is presently available at PPPL will be used, and this is listed also in Table D-1. Since current profile measurements are particularly important, special emphasis will be placed on installing the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam.

The long-term objective is to upgrade and expand the diagnostic set for the study  of fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics will provide input for advanced plasma control systems. Possible upgrade diagnostics that currently exist at PPPL include a “DensPak” array of Mirnov coils, a two-dimensional soft X-ray imaging system, beam emission spectroscopy, amplitude modulation reflectometry (in collaboration with the Oak Ridge National Laboratory and the EURATOM/CIEMAT group in Madrid, Spain), a multichannel fast neutral analyzer, and fast and slow neutron counters. New concepts and systems developed by the national NSTX team to implement these goals also will be encouraged. For this reason, the component design features and system performance for Day One and baseline diagnostics only are described in this document.

Table D-1: Day One* and Baseline Diagnostics Summary

| System | Function |
| --- | --- |

| Plasma current Rogowski coils* | Total plasma current |
| --- | --- |

| Eddy current Rogowski coils* | Vertical position control |
| --- | --- |

| Voltage loops* | Poloidal flux for plasma control |
| --- | --- |

| Bq, Br coils* | Magnetic fluctuations |
| --- | --- |

| Mirnov coils* | Plasma control/magnetic fluctuations |
| --- | --- |

| Flux loops* | Toroidal flux for plasma control |
| --- | --- |

| Segmented loops* | Flux measurements on passive plates |
| --- | --- |

| Visible TV camera* | External shape for plasma control |
| --- | --- |

| IR camera* | Heat loads |
| --- | --- |

| Residual gas analyzers* | In-vessel gases |
| --- | --- |

| Multichannel TVTS1 (single pulse) | Electron temperature and density |
| --- | --- |

| CHERS2 (with neutral probe beam) | Ion temperature and toroidal rotation |
| --- | --- |

| MSE3 (with neutral probe beam) | Plasma current density profile |
| --- | --- |

| Multichannel bolometer | Radiated power profile |
| --- | --- |

| Visible continuum array | Zeff(r) |
| --- | --- |

| 1 mm microwave interferometer | Line-integrated plasma density |
| --- | --- |

| Survey spectrometer (SPRED) | Plasma impurities |
| --- | --- |

| Hα detectors | Edge recycling |
| --- | --- |

| Langmuir probes | Edge parameters |
| --- | --- |

| Soft X-ray imaging system | Plasma instabilities and fluctuations |
| --- | --- |

The abbreviations used in the table are as follows.

1TVTS: Multichannel Thomson Scattering System

2CHERS: Charge-Exchange Recombination Spectroscopy

3MSE: Motional Stark Effect Polarimetry

WBS 4X -  Diagnostics Design Description

1.0 Design Description

1.1 System Design Features

1.1.1 Functions

1.1.1.1 Day One Diagnostics

The function of Day One diagnostics is to support basic machine operations. They are listed in Table D-1.

1.1.1.2 Baseline Diagnostics

The function of the baseline diagnostics is to support the near-term emphasis of the NSTX program. This is effected by detailed measurements of plasma profiles for understanding the basic confinement and transport properties of short pulse, Ohmically-heated NSTX discharges. The baseline diagnostics are listed in Table D-1.

1.1.1.3 Upgrade Diagnostics

The function of the upgrade diagnostics is to support the long-term NSTX program objective of studying fluctuations and transport in long pulse, auxiliary-heated NSTX plasmas. In addition, plasma diagnostics will provide input for advanced plasma control systems. Possible upgrade diagnostics are given in Table D-1.1.1.3-1.

Table D-1.1.1.3-1: Upgrade NSTX Diagnostics

| System | Function |
| --- | --- |

| “DensPak” Mirnov Coils | Mode identification/magnetic fluctuations |
| --- | --- |

| Diamagnetic Loop | Stored energy |
| --- | --- |

| Soft X-Ray Imaging System | Tomographic reconstruction of plasmas |
| --- | --- |

| AM Reflectometry | Density profiles/fluctuations |
| --- | --- |

| Beam Emission Spectroscopy | Density fluctuations |
| --- | --- |

| Multichannel Radiometer | Fluctuations/“ballooning modes” |
| --- | --- |

| Fast Scanning Edge Probe | Fluctuations/edge plasma profiles |
| --- | --- |

| Multichannel Charge-Exchange Analyzer | Fast ion distributions |
| --- | --- |

| Gas Neutron Counters | Thermal neutrons |
| --- | --- |

| Plastic Scintillator Detectors | Fast neutron fluctuations |
| --- | --- |

a. Divertor probe arrays and Rogowski coils on backing plate supports will be installed as upgrade diagnostics.

b. The national NSTX team will be encouraged to develop new diagnostic concepts and systems to implement the program objectives of detailed physics understanding of high-performance plasmas and advanced plasma control.

1.1.2 Configuration Description and Essential Features

1.1.2.1 Internal Diagnostics

a. All coils, loops, and other diagnostics mounted inside the vacuum vessel are to be compatible with a 150o C vacuum vessel bakeout temperature.

b. All coils, loops, and other diagnostics mounted inside the vacuum vessel are to be installed in a way that minimizes plasma damage. This includes protection against high voltages and large heat loads, particularly for components located on the center column.

c. Thermocouples are to be placed on the backing plates of the plasma facing tiles in the main plasma chamber (including the passive stabilizer plates) and the divertor region.

1.1.2.2 Diagnostic Vacuum System

a. The vacuum system, if required, for each diagnostic are to be included in the diagnostic fabrication task. This includes, for example, forepumps, turbopumps, cryopumps, and associated controllers.

b. The exhaust outlets for forepumps are to be provided as part of the NSTX vacuum pumping system (WBS 31).

1.1.2.3 Electrical Power and Isolation

a. All diagnostics are to have electrical isolation from the vacuum vessel that is rated for (TBD) kV.

b. All connections to and from diagnostic equipment in the NSTX test cell are to be with optical fiber. All connections between the data acquisition systems for each diagnostic are to be with optical fiber.

c. All diagnostics and associated data acquisition systems in the NSTX test cell are to receive AC input power through isolation transformers rated for a (TBD) kV DC continuous test, (TBD) kV DC rms one minute hi-pot test.

d. There will be two independent AC power distribution systems. One 60 Hz, 120 V up to 480 V system will be designated for diagnostic signal processing, data acquisition, and control subsystems. The other 60 Hz, 120 V up to 480 V system will be used to power motors and other “noisy” or high power equipment.

e. Each diagnostic will be independently grounded to a single point ground as determined by the NSTX machine grounding specifications.

1.2 Layout

The locations of the diagnostics on NSTX are shown in Fig. D-1.2-1.

Schematic Plan View of NSTX Showing Location of Diagnostics

Figure D-1.2-1

1.3 Component Design Features & System Performance for Day 1 & Baseline Diagnostics

1.3.1 General Magnetic Diagnostics

1.3.1.1 Design features

Description (including physical characteristics):

loops and coils at numerous toroidal and poloidal locations:

Plasma current Rogowski coils: up to 2 toroidal locations.

Eddy current Rogowski coils: up to 27 on passive stabilizers.

Voltage loops: up to 9 poloidal locations.

Bθ, Bρ coils: up to 20 locations.

Mirnov coils: up to 120 locations.

Flux loops: up to 20 locations.

Segmented loops: up to 10 locations.

Fig. 1.3.1-1 shows the locations of the flux loops (‘x’) and Mirnov coils (‘•’) in a poloidal cross section of NSTX above the horizontal midplane. The same arrangement of flux loops and Mirnov coils will occur below the horizontal midplane. Measurements of the toroidal field coil voltage and current will be used to obtain the plasma kinetic energy from the diamagnetic effect. The installation concept is illustrated schematically in Fig. 1.3.1-2.

The flux loops are toroidally continuous, and are located at 24 discrete poloidal locations. They are placed in grooves in the microtherm insulation on the side facing the OH solenoid. Their twisted leads will go through a groove in the microtherm to a vacuum feedthrough.

The Mirnov coils are located at 24 discrete poloidal at two discrete toroidal locations. There will be Mirnov coils in the horizontal midplane at 10 additional toroidal locations on the outer major radius side and 10 additional locations on the inside major radius side of the vacuum vessel. Their dimensions are (TBD) in order to fit inside cavities in the carbon tiles lining the inside of the vacuum vessel. They will also have a (TBD) slit in front of them to enable high frequency response. Their twisted leads will go through a channel in the tiles to a vacuum feedthrough. The coils, leads, and mounts must be compatible with the 350oC bakeout temperature of the NSTX vacuum vessel.

Parameters measured (T, q, n, etc.):

equilibrium plasma configuration, plasma control parameters, MHD fluctuations

1.3.1.2 System Performance

Resolution (spatial, time, etc.): 10 µsec time resolution.

Measurements limits (such as density range): digitization limit (coils can sense frequencies in the 200-300 kHz range).

Method of calibration:

Mirnov coils and Rogowski loops calibrated with known currents.

Amplifiers usually left on during normal operations.

No special emergency shutdown procedure.  Diagnostic is passive system mounted inside NSTX vacuum vessel.

Schematic Elevation of NSTX Showing Location of Magnetic Diagnostics

Figure D-1.3.1-1

Schematic Concept for Mounting Flux Loops and Mirnov Coils

Figure D-1.3.1-2

1.3.2 Plasma TV Camera

1.3.2.1 Design Features

Description (including physical characteristics):

tangential views of main plasma chamber and divertor regions.

Parameters measured (T, q, n, etc.):

plasma shape, edge recycling.

Plasma coverage (edge, edge to middle, # of channels, etc.):

imaging of edge region.

1.3.2.2 System Performance

Resolution (spatial, time, etc.):

20 msec temporal resolution, 5 cm spatial resolution.

Detector voltage control conducted remotely through CAMAC interface.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

1.3.3 IR Camera

1.3.3.1 Design Features

Description (including physical characteristics):

tangential views of main plasma chamber and divertor regions.

Parameters measured (T, q, n, etc.):

temperature waveform, incident power densities, plasma-wall interactions.

Plasma coverage (edge, edge to middle, # of channels, etc.):

top and bottom divertor regions, all passive stabilizer surfaces.

1.3.3.2 System Performance

Resolution (spatial, time, etc.):

20 msec temporal resolution, 5 cm spatial resolution, 14o x 18o field-of-view.

Measurements limits (such as density range):

-20oC to 1500oC temperature range, 8 to 12 micron spectral range, requires ZnSe (or equivalent) vacuum windows.

Method of calibration:

calibrated offline using calibrated black body source.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

1.3.4 Thomson Scattering (TVTS)

1.3.4.1 Design Features

Description (including physical characteristics): single pulse, multi-point.

Parameters measured (T, q, n, etc.):

Electron temperature and density profiles

Plasma coverage (edge, edge to middle, # of channels, etc.):

55 spatial points radially on horizontal midplane.

1.3.4.1 System Performance

Resolution (spatial, time, etc.):

1.2 cm spatial resolution at single time point per shot.

Measurements limits (such as density range):

50 eV to 5 keV.

Method of calibration:

in situ (using Rayleigh scattering in gas-filled torus), white plate source traceable to NBS standard.

Laser firing and detector voltage control conducted remotely through CAMAC interface.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

Under routine operation in support of NSTX experiments, the laser, its beam path, and target area are completely enclosed.  In addition, the laser room and the laser shack on the NSTX platform are kept locked.

1.3.5 Charge-Exchange Recombination Spectroscopy (CHERS)

1.3.5.1 Design Features

Description (including physical characteristics):

measure Doppler-broadened line profiles of transitions excited by charge-exchange recombination reactions between fast neutrals from heating beam and fully ionized low-Z impurities.

Parameters measured (T, q, n, etc.):

Ion temperature and toroidal rotation velocity profiles

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge and core, 12 channels.

1.3.5.2 System Performance

Resolution (spatial, time, etc.):

3 msec time resolution; 3 cm spatial resolution.

Measurements limits (such as density range):

core channels have density limit in excess of <ne> = 7 x 1013 cm-3 or nel greater than 4 x 1015 cm-2, requires use of diagnostic neutral probe beam at energy in excess of 40 keV.

Method of calibration:

use Hg lamp to calibrate dispersion (vs. wavelength) and use Hg lines to calibrate experimental line width (vs. wavelength).

Image intensifier and detector voltages set remotely through CAMAC interface.  Requires operation of single heating beam (northwest) for neutral doping.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

1.3.6 Motional Stark Effect Polarimetry (MSE)

1.3.6.1 Design Features

Description (including physical characteristics):

measures magnetic field pitch angle from polarized emission from neutral probe beam using photoelastic modulators and photomultiplier detectors.

Parameters measured (T, q, n, etc.):

q.

Plasma coverage (edge, edge to middle, # of channels, etc.):

core and edge, 10 spatial channels, upgradeable to 18 channels.

1.3.6.2 System Performance

Resolution (spatial, time, etc.):

3 cm spatial resolution, 10 msec temporal resolution.

Method of calibration:

optical alignment by backlighting fiber bundles and recording images on target when vacuum vessel is vented to atmosphere.

AC power and modulator voltage controlled remotely through CAMAC interface.  All other voltages controlled locally using power supplies behind NSTX shielding wall. Requires operation of diagnostic neutral probe beam in hydrogen.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

1.3.7 Bolometry

1.3.7.1 Design Features

Description (including physical characteristics): tangential array in horizontal midplane, two divertor views.

Parameters measured (T, q, n, etc.):

Radiated power.

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge and core, 15 channels in tangential array.

1.3.7.2 System Performance

Resolution (spatial, time, etc.):

4 cm spatial resolution, 5 msec time resolution.

Method of calibration:

electrical heat pulse.

No special emergency shutdown procedure.  Diagnostic does not have separate vacuum system independent of NSTX.

1.3.8 Visible Bremsstrahlung Array

1.3.8.1 Design Features

Description (including physical characteristics): multi-point array.

Parameters measured (T, q, n, etc.):

Zeff

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge & core, 15 radial points.

1.3.8.2 System Performance

Resolution (spatial, time, etc.):

4 cm spatial resolution, 1 msec time resolution.

Measurements limits (such as density range):

Method of calibration:

tungsten lamp.

Detector voltage control conducted through CAMAC interface.  Quartz shutters remotely controllable from control room.  Fibers installed over radiation shielding wall between NSTX platform and auxiliary control room.

No special emergency shutdown procedure.  Diagnostic is optical system which is exterior to NSTX vacuum vessel.

1.3.9 Microwave Interferometer

1.3.9.1 Design Features

Description (including physical characteristics):

vertical chord through center of plasma.

Parameters measured (T, q, n, etc.):

Line-integrated plasma density.

Plasma coverage (edge, edge to middle, # of channels, etc.):

single 2 mm channel.

1.3.9.2 System Performance

Resolution (spatial, time, etc.):

1 µsec time resolution.

Measurements limits (such as density range):

limited by digitization rate.

Method of calibration:

use calibrated rf generator.

Detector voltage adjusted on power supply on NSTX platform prior to operations.

No special emergency shutdown procedure.  Diagnostic does not have vacuum system contiguous with NSTX vacuum vessel.

1.3.10 Survey, Poor Resolution, Extended Domain (SPRED) VUV Spectrometer

1.3.10.1 Design Features

Description (including physical characteristics): vacuum ultraviolet spectrometer with intensified self-scanning photodiode array detector.

Parameters measured (T, q, n, etc.):

impurity distributions.

Plasma coverage (edge, edge to middle, # of channels, etc.):

core or edge region (depending on presence or absence of view across neutral beam), 1024 pixels.

1.3.10.2 System Performance

Resolution (spatial, time, etc.):

5 cm spatial resolution, 1 msec time resolution, ≈2Å spectral resolution.

Measurements limits (such as density range):

100Å-1200Å.

Method of calibration:

mechanical alignment of grating with He-Ne laser and He discharge for optical alignment.

Detector voltage set remotely through CAMAC interface.  Machine valve is interlocked with NSTX pressure and SPRED pressure.

No special emergency shutdown procedure.  All valves will shut and voltage will be automatically turned off in event of vacuum failure or electrical power failure.

1.3.11 Hα Detectors

1.3.11.1 Design Features

Description (including physical characteristics):

Diode detectors with interference filters at various toroidal locations for edge Hα and Dα radiation measurements as a function of space and time.

Parameters measured (T, q, n, etc.):

edge Hα and Dα densities

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge

1.3.11.2 System Performance

Resolution (spatial, time, etc.):

1 msec time resolution.

Measurements limits (such as density range):

edge region of plasma.

1.3.12 Langmuir Probes

1.3.12.1 Design Features

Description (including physical characteristics):

Langmuir probes at various toroidal locations for edge parameter measurements as a function of space and time.

Parameters measured (T, q, n, etc.):

edge potential, Te, and ne

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge

1.3.12.2 System Performance

Resolution (spatial, time, etc.):

0.004 msec time resolution.

Measurements limits (such as density range):

probe access restricted to edge region of plasma.

1.3.13 Soft X-Ray Diode Imaging System (XIS)

1.3.13.1 Design Features

Description (including physical characteristics):

detect soft X-ray emission with vertically-viewing and horizontally-viewing diode arrays.

Parameters measured (T, q, n, etc.):

MHD fluctuations

Plasma coverage (edge, edge to middle, # of channels, etc.):

edge and core, 32 channels along each view.

Special features:

fluctuation diagnostic, can be used in conjunction with plasma shape diagnostics for tomographic reconstructions

1.3.13.2 System Performance

Resolution (spatial, time, etc.):

2-3 cm spatial resolution, 0.001 msec temporal resolution.

Measurements limits (such as density range):

maximum resolvable mode number is m = 2, modes up to m = 6 distinguishable but not resolvable with 32 channels.

Method of calibration:

diodes calibrated off-line with X-ray source.

Detector voltage control conducted remotely using VAX computer through CAMAC interface.

No special emergency shutdown procedure.  Diagnostic is passive system mounted inside NSTX vacuum vessel.

2.0 System Performance

Included in Section 1.3.

3.0 Interfaces

3.1 Diagnostic Port Utilization

a. The plasma diagnostics on NSTX will connect to the vacuum vessel and support structure (WBS 12) at the machine interface vacuum valve and/or the points of mechanical support for the diagnostic inside the vacuum vessel. Four large midplane ports and at least three dome ports will be available for diagnostics. The port utilization will be as given in Table D-3.1-1.

Table D-3.1-1: NSTX Diagnostic Port Requirements

| Location | Diagnostic System |
| --- | --- |

| Large midplane port Dome port above midplane for XIS Dome port below midplane for TVTS | Multichannel TVTS (single pulse) Soft X-ray imaging system (XIS) Survey spectrometer (SPRED) |
| --- | --- |

| Large midplane port | Multichannel bolometer Single channel bolometers Visible continuum array |
| --- | --- |

| Large midplane port | Visible TV camera IR camera Hα detectors |
| --- | --- |

| Large midplane port | CHERS MSE 1 mm microwave interferometer Langmuir probes |
| --- | --- |

| Large midplane port (for neutral heating beam during NSTX upgrade phase) | Diagnostic neutral probe beam (will be removed during NSTX upgrade phase) |
| --- | --- |

| Dome port | Residual gas analyzer |
| --- | --- |

b. The neutral probe beam will use a large midplane port that is reserved for a neutral heating beam during the NSTX upgrade phase. The neutral heating beam will perform the function of the neutral probe beam during the NSTX upgrade phase.

c. The visible TV and infrared (IR) cameras will have tangential views of main plasma chamber and divertor regions

d. The multichannel bolometer will consist of a tangential array in horizontal midplane.

e. The two single channel bolometers will have views of the upper and lower divertors.

f. The dumps for the TVTS laser and the neutral probe beam will be located between ports.

g. All connections between the diagnostics and the vacuum vessel will be compatible with a 150o C vacuum vessel bakeout temperature.

3.2 Data Acquisition System

a. Planning and costing for each diagnostic will include the device control and data acquisition modules that interface with the data highway of the NSTX Data Acquisition System (WBS 62).

b. The diagnostics data acquisition summary is given in Table D-3.2-1. To establish requirements for Central I&C, each Data Acquisition signal is included in one of the following general categories:

Low Speed ==> Data Acquisition signals requiring sample rates >100 Hz and ≤2 KHz.

Moderate Speed ==> Data Acquisition signals  requiring sample rates >2 KHz and ≤100 KHz.

High Speed ==> Data Acquisition signals  requiring sample rates >100 KHz and ≤1 MHz.

The Average Sampling Time for each signal category (per NSTX pulse) is also included in the table below.

Table D-3.2-1: Data Acquisition Summary

|  | Number of Signals & Average Sampling Time per Pulse |
| --- | --- |

| NSTX Diagnostics | Low Speed >100Hz & ≤2KHz | AVG SMPL Time (sec/ pulse) | Mod. Speed >2KHz & ≤100 KHz | AVG SMPL Time (sec/ pulse) | High Speed >100 KHz & ≤1MHz | AVG SMPL Time (sec/ pulse) |
| --- | --- | --- | --- | --- | --- | --- |

| Day One Diagnostics | 250 | 1 |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |

| Baseline Diagnostics |  |  | 150 | 1 |  |  |
| --- | --- | --- | --- | --- | --- | --- |

| Upgrade Diagnostics |  |  | 50 | 5 | 100 | 5 |
| --- | --- | --- | --- | --- | --- | --- |

| TOTALS | 250 |  | 200 |  | 100 |  |
| --- | --- | --- | --- | --- | --- | --- |

c. The existing PPPL computer network will be used for diagnostic data processing. The present bus architecture connects alpha workstations, X-terminals, etc. in a network that includes collaboration links. The network supports distributed diagnostic control and data acquisition and centralized data management.

d. The interface between the central I&C data highway and the diagnostic device control and data acquisition modules will be a hybrid that can accommodate existing CAMAC equipment as well as hardware based on newer standards.

4.0 Design Basis

4.1 Standards & Codes

Not applicable.

4.2 Supporting Analysis

Not applicable.

5.0 Design Implementation

5.1 Prototypes

Not applicable.

5.2 Mock-ups

Not applicable.

5.3 Component Procurement and Fabrication

Not applicable.

5.4 System Assembly, Installation, and Testing

Not applicable.

A.1 Special NSTX Diagnostic Needs

A.1.1  Current Profile Measurements

Because of the importance of current profile measurements, installation of the motional Stark effect diagnostic and its dedicated diagnostic neutral probe beam has high priority. The special issues that must be addressed are as follows.

•	Diagnostic Neutral Probe Beam (NPB):

1) Operational reliability - The NPB was commissioned at the Culham Laboratory (UK) for operation with 100 ms pulses at 80 kV. This was never achieved at PPPL because of power supply difficulties. They should be corrected as part of the task of moving the beamline to the NSTX platform (test cell) and installing a new control console in the NSTX control room.

2) Long pulse operation - The existing power supplies (when working reliably) permit 1 s operation. New grids are needed to accommodate this pulse length. Longer (5 s) NSTX discharges can then be accommodated by operating the NPB in pulsed mode (e. g., 20 ms pulses every 100 ms for 50 time points in 5 s) but major source and power supply upgrades are needed for 5 s DC operation.

3) Availability of the neutral probe beam - PBX-M will require the diagnostic neutral probe beam (NPB) for at least a year of operations, while a new scheme to instrument one of the heating beams for motional Stark effect (MSE) polarimetry is being tested. The use of a heating beam on PBX-M is speculative because it has a lower injection energy than the NPB, so we need to determine if the improved signal from the higher heating beam current will compensate for it. If PBX-M resumes operation in FY99, the schedule shall be to move the NPB to NSTX at the beginning of FY2000 at the earliest, and begin current profile measurements in the middle of that fiscal year. This will avoid having to develop MSE diagnostics simultaneously on NSTX (see below) and PBX-M, which are both really new because of the novel physics constraints.

•	Motional Stark Effect (MSE) Diagnostic:

1) Feasibility of measurement - The magnitude of the NSTX toroidal field relative to the NPB injection energy and the NPB trajectory relative to the MSE sightlines must be evaluated to determine if the Stark-shifted line can be resolved. Low field, low aspect ratio, and port access are all potential difficulties.

A.1.2  Electron Temperature Profile Measurements

Electron temperature and density profiles will be critical to the understanding of the confinement and stability of NSTX plasmas. The initial phase of NSTX operations will be characterized by short pulses and high repetition rates. This is similar to the present situation on PBX-M, where the temporal evolution of electron temperature and density profiles is determined from multiple discharges using a TV Thomson scattering (TVTS) system with a single-pulse laser. In the long term, however, it is desirable to measure the electron temperature and density profiles as a function of time in a single discharge. This is particularly true when the NSTX pulse lengths are extended to 5 s, and profile information is also needed for feedback applications. Two options for a multi-time-point TVTS system are summarized as follows.

1) Install multipulse glass laser and upgrade imaging and computer systems at a cost of ≈ $0.5M.

2) Install completely new Nd:YAG laser-based system at a cost of ≈ $1.5M

The first option is attractive because of its lower cost, but the overall diagnostic concept has not been tested. The cost for the second possibility is much higher, but the system has been demonstrated for profile measurements in multi-second discharges on the DIII-D tokamak. As a system in the upgrade set, the choice of which TVTS option to pursue will depend on when the NSTX program needs the diagnostic and the funds available for it.