National Spherical Torus Experiment

NSTX

SPECIFICATION

FOR

TF COIL OUTER LEG ASSEMBLY

NSTX-SPEC-13-043

Revision 0

DATED: November 12, 1997

PREPARED BY:___________________________________

J.H. Chrzanowski

NSTX Tokamak Systems WBS Manager

REVIEWED BY:___________________________________

Charles Neumeyer

NSTX  Engineering Manager

APPROVED BY:___________________________________

P.J. Heitzenroeder

Deputy NSTX Project Manager

PRINCETON UNIVERSITY

PLASMA PHYSICS LABORATORY

(PPPL)

P.O. BOX 451

PRINCETON, N.J. 08543

RECORD OF CHANGE

| Revision | Date | TRB | Description of Change |
| --- | --- | --- | --- |

TABLE OF CONTENTS

1.0	SCOPE:  ...................................................	1

2.0	APPLICABLE DOCUMENTS: ...........................	1

2.1	DRAWINGS: ............................................................	1

2.2	STANDARDS AND CODES: .........................................	1

2.3	GENERAL STATEMENT: ............................................	2

3.0 	REQUIREMENTS: ........................................	2

3.1	EQUIPMENT DEFINITION: .........................................	2

3.1.1	General Design Features: ...............................................	2

3.1.2	Materials: ..................................................................	2

3.1.2.1	Copper Conductor: ..............................................	2

3.1.2.2	Insulation System: ...............................................	3

3.1.2.3	Soldering and Braze Materials:..................................	4

3.1.2.4	Degreasing/Cleaning Solvent:.................................	4

3.1.2.5	Machined Coil Lead Blocks and Electrical Flags: ............	4

3.1.2.6	Grit Blasting Materials: .........................................	4

3.1.2.7	Copper Primer: ...................................................	5

3.1.2.8	Conductor Coolant Tube ........................................	5

3.1.2.9	Press-Mold Release Agent ......................................	5

3.2	COIL  FABRICATION ..................................................	5

3.2.1	Copper Conductor Delivery: .............................................	5

3.2.2	Conductor Inspection: ....................................................	5

3.2.3	Conductor Forming: ......................................................	5

3.2.4	Conductor Inspection: ....................................................	5

3.2.5	Electrical Flag Installation: ...............................................	6

3.2.6	Coolant Tube Installation: ................................................	6

3.2.7	Flow and Hydrostat Tests: ................................................	6

3.2.8	Conductor Cleaning: ......................................................	6

3.2.9	Material Protection: .......................................................	6

3.2.10	Conductor Grit Blasting: .................................................	7

3.2.11	Cleaning After Grit Blasting: ............................................	7

3.2.12	Primer Application: .......................................................	7

3.2.13	Turn Insulation: ...........................................................	7

3.2.14	Lead Area Turn Insulation: ..............................................	7

3.2.15	Assy. Of Outer TF Leg Conductors: ...................................	7

3.2.16	Outer TF Leg Assembly Test-Pre-Cure: ...............................	8

3.2.17	Outer TF Leg Turn Insulation Cure: ....................................	8

3.2.18	Outer Ground Wall Insulation: ..........................................	8

3.2.19	Outer TF Leg Groundwall  Cure: .......................................	8

3.2.20	Evaluation of Finished Surfaces: .......................................	9

3.2.21	Installation of TF Coolant Tube Supports: ............................	10

3.2.22	Final Inspection and Test: ...............................................	10

4.0	TEST REQUIREMENTS: ................................	10

4.1	RECEIVING INSPECTION-OUTER TF  CONDUCTOR .......	10

4.1.1	Dimensional Measurements/Inspections: ............................	10

4.2	IN-PROCESS TESTING: ............................................	10

4.2.1	Tests After Conductors have been Formed: .........................	10

TABLE OF CONTENTS

4.2.2	Tests After Electrical Flag Installation: ..............................	10

4.2.3	Tests After Coolant Tube Installation: ...............................	10

4.2.4	Tests After TF Conductor Assy. and Prior to Oven Cure: ........	11

4.3	FINAL ACCEPTANCE TESTS: .....................................	11

4.3.1	Inspection: ...............................................................	11

4.3.2	Coil Flow/Pressure Tests: .............................................	12

4.3.3	Final Electrical Tests: ...................................................	12

5.0	QUALITY ASSURANCE REQUIREMENTS: ..........	12

5.1	INSPECTION/ SURVEILLANCE/ AUDIT BY PRINCETON: ..	12

5.2	SUBCONTRACTOR’S RESPONSIBILITY

FOR CONFORMANCE: ..............................................	12

5.3	QUALITY ASSURANCE PLAN: ....................................	13

5.4	MANUFACTURING, INSPECTION AND

TEST PLAN (MIT): ...................................................	13

5.5	INSPECTION AND TEST PROCEDURES: .......................	13

5.6	DOCUMENT TRACEABILITY AND RECORDS: ...............	13

5.7	EQUIPMENT/MATERIAL IDENTIFICATION

AND STATUS: .........................................................	13

5.8	CALIBRATION OF TEST AND MEASURING

EQUIPMENT: ..........................................................	13

5.9	CONTROL OF SPECIAL PROCESSES ...........................	14

5.10	NONCONFORMANCES: ............................................	14

5.11	RELEASE FOR SHIPMENT FORM: ...............................	14

5.12	PROCESS HISTORY: ................................................	14

5.12.1	Material Certifications: ........................................	14

5.12.2	Inspection and Test Reports: ................................	14

5.12.3	Fabrication Drawings: .........................................	14

5.12.4	Certificate of Compliance (C of C): .........................	15

5.12.5	Nonconformance Reports: ...................................	15

5.13	PPPL RECEIVING/INSPECTION .................................	15

6.0	DOCUMENTATION REQUIREMENTS: ...............	15

6.1	Proposal Technical Documentation: .................................	15

6.2	Documentation After Contract Award and Prior to Fabrication:..	15

6.3	Documentation Prior to Shipment: ...................................	16

7.0	SHIPPING STORAGE AND HANDLING: ...........	16

8.0	DELIVERABLES: .......................................	16

9.0	ATTACHMENTS:

APPENDIX I	Outer TF Leg Assembly drawings: ..........	17

1.0	SCOPE:

a. This specification defines the manufacturing requirements for the National Spherical Torus Experiment (NSTX) Toroidal Field (TF) Outer Leg Assemblies which contain the outer legs of the TF coil system.

b. Subcontractor shall furnish all material, labor  and facilities necessary to complete the work as defined in the contract documents subject to the qualifications, conditions, and exceptions noted herein.

2.0	APPLICABLE DOCUMENTS:

2.1	DRAWINGS

a. 	Drawings identified in Appendix I describe the equipment to be provided under this specification.

b. 	Dimensions in drawings and specifications refer to conditions at room temperature (20° C/ 68° F).

2.2	STANDARDS AND CODES:

a. 	Materials and manufacturing/test methods used in fabrication of the equipment covered by this specification shall comply with the latest revision, in effect at date of purchase order, of the following currently approved applicable regulations, safety codes, specifications and standards, including applicable technical definitions as acknowledged and accepted in industry.

ASTM American Society for Testing and Materials

ASTM B260-62T (BCuP-5), Brazing Alloys

ASTM B32-Sn95, Soft Solder

CDA Copper Development Association

CDA Standards Handbook, Part 2 (Wrought Products)  Alloy Data,  CDA Alloy No. 10500

Conductor:

Per PPPL specification NSTX-SPEC-13-038-0 “TF Outer Leg Copper Conductor”.

IEEE- Institute of Electrical and Electronic Engineers

IEEE #4, Techniques for Dielectric Tests

NEMA-National Electrical Manufacturers Association

Grades and specifications for materials developed for the electric and electronics industries.

b. 	The above Standards and Codes set forth the minimum requirements.  They may be exceeded by Subcontractor with written permission from PPPL, if in Subcontractor’s judgment, superior or more economical designs or materials are available for successful and continuous operations, as required by the specification.

2.3	GENERAL STATEMENT

Subcontractor agrees, represents and warrants that all services, designs, equipment and material sold or otherwise provided to PPPL by Subcontractor comply with applicable, Federal, State and Local laws, regulations and codes, and all applicable specifications and standards including those specified above, in each case as in effect at the date of order placement. Subcontractor shall provide PPPL with a Certificate of Compliance identifying the codes, standards and regulations complied with, and any exceptions taken,  in the design and fabrication of deliverable items.

3.0 	REQUIREMENTS:

3.1 EQUIPMENT DEFINITION

3.1.1 	General Design Features

a. 	Dimensional characteristics of the coils are given in the drawings identified in Appendix I.

b. 	Outer TF Leg Assembly insulation builds are given in the drawings identified in Appendix I.

c. Tolerances are indicated on the drawings identified in Appendix I. Strict adherence to these dimensions is critical to the assembly and performance of the subject equipment.

3.1.2	Materials:

3.1.2.1 	Copper Conductor

a. 	The conductor will be CDA alloy 10500, silver bearing copper, (10 troy ounces per ton 0.034% min.) with Yield Strength of 30 ksi (0.5% elongation) and Ultimate Tensile Strength 38 ksi minimum (15% elongation).

b. 	The copper conductor will be provided by PPPL.  All generated scrap and excess material furnished by PPPL shall remain the property of PPPL. A material inventory, defining quality, condition and quantity of excess material shall be maintained and submitted at the conclusion of the job. PPPL shall instruct Subcontractor as to the disposition of excess materials at PPPL’s expense.

c. 	The conductors will be supplied by PPPL to the Subcontractor with the coolant tube grooves already extruded in each conductor.

3.1.2.2	Insulation System

a. 	The turn insulation shall consist of (3) half-lapped layers of B-staged epoxy-glass tape (Vonroll Isola Product no. 2258XS).

b. 	The B-staged epoxy-glass turn insulation shall be provided by PPPL. All remaining insulating materials required to manufacture the Outer TF Leg assemblies including outer groundwall shall be provided by the Subcontractor.  All excess material furnished by PPPL shall remain the property of PPPL. A material inventory, defining quality, condition and quantity of excess material shall be maintained and submitted at the conclusion of the job.  PPPL shall instruct Subcontractor as to the disposition of excess materials at PPPL’s expense.

Note:  The Vonroll Isola products must be kept in a refrigerator or freezer to maintain shelf life. ( ≤ 40° F)

c.	The outer ground insulation shall be “Scotchply” type No. 1003 B-staged epoxy-glass tape.   The final radial build after final cure shall be in accordance with the drawings identified in Appendix I.

e.	The dielectric shields in the joint areas shall be Kapton polyimide film (0.005  in. thk.) in sheet form without adhesive backing.

f. 	Details concerning the insulating materials are as follows:

Turn to Turn Insulation (Vonroll Isola Product no. 2258XS with interliner)

Epoxy-glass pre-preg insulation in tape form

Resin Type:	TGDM epoxy, aromatic amine cure agent

Composite Technology Development (CTD), Boulder, Co

Glass tape:  	S-2 ( satin weave) std. silane finish

Temperature Class:   180° C

Nominal Thk.: 	0.006 in. (0.152 mm.) tape

Nominal Width:  	1.0 in. tape

Groundwall Insulation (3M-Scotchply #1003)

Epoxy-glass pre-preg insulation in tape form

Glass Tape:  Type “K” filament

Nominal Thickness: 0.010 in. (0.254 mm)

Nominal Width:  2.0 in. Tape

Kapton Sheet- “Dupont” product

Type H polyimide film

Nominal Thickness:  0.005 in.

Nominal width:  Sheet form

Temperature Class:	180° C

3.1.2.3	Soldering and Braze Materials:

a.	The electrical flags shall be brazed to the coil leads using Sil-Fos braze material a product of “Handy & Harmon”. No flux shall be used with the Sil-Fos material.

Composition: 15% Ag, 80% Cu, 5% P

Melting Point  1185° F (640° C)

Flow Point  1130° F (705° C)

b.	Certification of the brazing alloys shall be in accordance with ASTM specifications noted in section 2.2, and shall be submitted as part of the document package.

c.	The copper coolant tube shall be soft-soldered into the TF conductor using 95%/5% Tin-Antimony soft-solder with a melt range of 453-464 °F.  Subcontractor shall recommend suitable liquid flux.  Any changes from this shall require PPPL approval.

d.	The Nibco fittings shall be soft-soldered onto the ends of the copper coolant tubes using 95%/5% Tin-Antimony soft-solder with a melt range of 453-464 °F.  Subcontractor shall recommend suitable liquid flux.  Any changes from this shall require PPPL approval.

3.1.2.4	Degreasing/Cleaning Solvent:

All conductors shall be degreased/cleaned using a solvent which has been pre-approved by PPPL.  This solvent must be able to dissolve grease, tar, wax, adhesives, oils and other soils, and be residue free.

3.1.2.5	Machined Coil Lead Blocks and Electrical Flags:

Machined blocks for coil leads shall be fabricated from copper bar or plate.  Alloy shall be CDA alloy no.10400 or 10700 oxygen free copper having a minimum yield strength of 30,000 psi (0.5% elong.).

3.1.2.6	Grit Blasting Materials:

Grit blasting shall be performed using aluminum oxide with a grain size of 80-120 mesh.

3.1.2.7	Copper Primer

The primer shall be Araldite DZ-80-1, manufactured by Ciba-Geigy (UK) Limited Plastics Division.  No substitutions permitted.

3.1.2.8	Conductor Coolant Tube

The coolant tube which shall be soldered into the TF conductors shall be seamless Air Conditioning and Refrigeration  (ACR) copper tubing with 0.500 in. outside diameter and 0.032 in. wall thickness.  The tubing shall be procured with the outside diameter pre-tinned prior to soldering into the conductor.

3.1.2.9	Press-Mold Release Agent

A suitable press-mold release agent shall be recommended by the subcontractor and approved by PPPL.

3.2	COIL  FABRICATION

This section describes a recommended procedure for fabricating the subject equipment.  The Subcontractor shall submit for approval by PPPL, a Manufacturing, Inspection, and Test (MIT) plan.  No deviations from the MIT plan will be accepted without PPPL approval.

3.2.1	Copper Conductor Delivery:

The copper conductors shall be delivered to the Subcontractor as straight length conductors.  The crates shall be inspected for visual damage by the Subcontractor and condition of crates documented.

3.2.2	Conductor Inspection:

Each copper conductor shall be removed from its container moved to an inspection station where they will be inspected per requirements in Section 4.0 paragraph 4.1.  An identification number shall be metal stamped onto the end of each conductor.

3.2.3	Conductor Forming:

The copper conductors shall be formed to the dimensions indicated on the drawings identified in Appendix I.  This work may be performed in house or by an outside vendor subcontracted by the coil manufacturer.  The bars must be cold formed without the use of external heat to minimize any changes to the yield and ultimate strengths of the extruded conductors

3.2.4	Conductor Inspection:

At the conclusion of the forming operations, the conductors shall be inspected per requirements in section 4.0 paragraph 4.2.1. Inspection and test results shall be recorded and included with the Process History package.

3.2.5	Electrical Flags Installation:

a.	The copper electrical connections shall be induction/torch brazed to the conductor ends following the forming of the conductors.   Braze materials are specified in section 3.1.2.3.  Braze process must be PPPL approved.  Braze inspection shall be in accordance with section 4.0 paragraph 4.2.2.

3.2.6 	Coolant Tube Installation

a.	The pre-tinned copper coolant tube shall be positioned and soft-soldered into the groove provided in each conductor.  The copper tube shall be continuous with no splices.

b.	Soft-solder the copper coolant tubes into the conductor grooves using  95%/5% Tin-Antimony soft-solder with a melt range of 453-464 °F.  Subcontractor shall recommend suitable liquid flux.

c.	Caution:  Extreme care must be taken to ensure that the heating process used for soldering does not effect the hardness of the copper extrusion.  Pre and post  Rockwell hardness measurements (per section 4.0) shall be taken by subcontractor to ensure compliance. This information shall be included in the Process History.

d.	Following the soldering operation, the excess solder shall be removed from the conductor so that the cross-section of the conductor conforms to the PPPL drawings.  Care should be taken so that the coolant tube is not damaged.  The conductor must also be cleaned to remove flux residue.

e. 	PPPL approved fittings shall then be installed onto the ends of each copper coolant tube using 95%/5% Tin-Antimony soft-solder.

3.2.7	Flow and Hydrostatic Tests:

Flow and hydrostatic tests of the soft-soldered coolant tube shall be performed per requirements in section 4.2.3 to verify integrity of the coolant tube and to provide a base line for future tests.

3.2.8 Conductor Cleaning

Each conductor shall be wiped down with degreasing solvent to remove excess oil, lubricant and grease.

3.2.9	Material Protection:

Cleaned copper, whether bare or insulated, shall be stored and processed in an environment free from metallic dust or other contaminants.  After cleaning, the copper surface and insulating materials shall be protected from skin oil, etc., by requiring shop personnel to handle conductors only while wearing clean, lint free, white cotton gloves.  Insulation and in-process subassemblies shall be processed and stored in controlled clean areas.

3.2.10 	Conductor Grit Blasting

Following the degreasing of the conductors and testing of the coolant tubes, the conductor shall be grit blasted to remove any surface oxidation and promote insulation bonding.  This process must be performed in a controlled  manner which will expose the conductor to grit blasting at a recommended rate of 12 seconds/ inch2. A pressure of 18-20 psi is recommended per nozzle at a distance of approximately 5-6 inches from the conductor.  It is essential that water and oil filters are located in the main air supply line to minimize the risk of introducing contaminants onto the conductors.

3.2.11	Cleaning After Grit Blasting

After grit blasting, the conductor shall be vacuum cleaned and wiped to remove all aluminum oxide grit remaining from the grit blasting.

3.2.12	Primer Application

A layer of DZ-80-1 primer will be applied by brushing or spraying. The final method for application and curing must be approved by PPPL.  It is important that the primer be applied with a thin uniform thickness.  Primer shall be air dried before proceeding.  The final cure of the primer shall occur during the insulation cure cycle.

3.2.13 	Turn Insulation

Three half-lapped layer of pre-impregnated (B-stage) epoxy-glass tape (Vonroll Isola Product no. 2258XS) shall be applied to the conductors following application of the primer.

Care shall be taken to avoid damage to the insulation in subsequent handling.  Damaged insulation shall be replaced with new insulation in accordance with a Subcontractor Prepared/ PPPL approved repair procedure which will maintain insulation continuity.

3.2.14 	Lead Area Turn Insulation

Turn insulation shall be tailored to fit neatly around water and electrical terminals.

3.2.15 Assembly of Outer TF Coil Leg Conductors

a. 	Insulated Outer TF conductors shall be assembled in accordance with approved drawings to form a Outer TF Leg Assembly. A vendor designed and supplied fixture (PPPL approved) is required to align the Outer Leg conductors during assembly. The electrical flags at each end of the bundle shall be utilized for the alignment.   During assembly, 0.005 inch thick Kapton dielectric shields shall be installed between conductors near the coil leads.

b.	Each coil assembly shall be given an ID number so that Process History of the assembly including ID of conductors used in each assembly can be maintained.

ID Example:  OTFL#1 .............OTFL#12

3.2.16	Outer TF Coil Leg Assembly Test- Pre-Cure

Each conductor within the Outer TF Leg Assembly shall be electrically tested against all other conductors shorted together and grounded per test requirements given in section 4.2.4.

3.2.17	Outer TF Leg Turn Insulation Cure

Note:  It is necessary to pre-cure the conductor turn insulation prior to applying the groundwall insulation.

a.	Two (2) half-lapped layers of 2 mil non-stick (Tedlar) tape shall be applied around the entire Outer TF Leg Assembly.

b. 	Two (2) half-lapped layers of shrinkable mylar tape shall be applied over the Tedlar layer.

c. 	Internal surfaces of a press mold (design and supplied by Subcontractor, approved by PPPL) shall be coated with pre-approved mold release agent.  The Outer TF Leg Assembly shall be installed into the press mold. The press mold shall be designed to provide sufficient compression of the B-stage insulating materials to insure that the resin flows into all voids and that the dimensions of the finished assembly are within the specified tolerances.

d. 	The coil bundle and mold shall be heated and cured to the following times and temperatures.

2 hours @ 177°C

7 hours @ 200°C

During this time the press mold shall apply sufficient compression (10% of nominal uncured B-stage insulation thickness) to the TF bundle to allow the Outer Leg Assembly to form a monolithic, void free structure with dimensions within the specified tolerances.

e. 	The subcontractor shall monitor and record cure temperature and time during the entire cure process.  The following temperatures shall be monitored and continuously recorded: the copper conductor, coil mold and air temperature. This information shall become part of the Process History per section 5.11.

f.  Following the insulation cure cycle, the Outer TF Leg Assembly shall be removed from the press-mold. The layers of Tedlar and shrinkable mylar shall be removed, and the assembly shall be inspected for dry spots, voids or other surface deformations.

3.2.18	Outer Ground Wall Insulation:

After removal from the mold and inspection, apply the Groundwrap insulation to the Outer TF Leg Assemblies.

a. 	Six (6) half-lapped layers of pre-impregnated B-stage tape (product #3M-Scotchply #1003) shall be applied over the pre-insulated conductors to provide the ground wall insulation.

b. 	All layers of ground insulation shall be applied in the same direction.

c. 	The ground wall build up shall result in achievement of the cured insulation thickness, specified in the coil drawings (Appendix I) after  compression during the final cure process.

d.	The ground insulation shall be tailored to fit neatly around the conductor end electrical flags.

3.2.19	Outer TF Leg Turn Groundwall Cure

a.	Two (2) half-lapped layers of 2 mil non-stick (Tedlar) tape shall be applied around the entire Outer TF Leg Assembly.

b. 	Two (2) half-lapped layers of shrinkable mylar tape shall be applied over the Tedlar layer.

c. 	The molding process used shall be developed by the vendor in accordance with the technical data supplied by the insulation Manufacturer.  The compression of the B-stage material shall be in the range of 14-20%.

d.	Internal surfaces of a press mold (design and supplied by Subcontractor, approved by PPPL) shall be coated with pre-approved mold release agent.  The Outer TF Leg Assembly shall be installed into the press mold. The press mold shall be designed to provide sufficient compression of the B-stage insulating materials to insure that the resin flows into all voids and that the dimensions of the finished assembly are within the specified tolerances.

e.	The cure cycle shall be as follows:

Rise.......	70°F - 295°F	2 ±1/4 hours

Hold ......	295°F ±10°F		5 ±1/4 hours

Cool ......	295°F - 70°F	1 1/2 ±1/4 hours

f. 	The subcontractor shall monitor and record cure temperature and time during the entire cure process.  The following temperatures shall be monitored and continuously recorded: the copper conductor, coil mold and air temperature. This information shall become part of the Process History per section 5.11.

g.  Following the insulation cure cycle, the Outer TF Leg Assembly shall be removed from the press-mold. The layers of Tedlar and shrinkable mylar shall be removed, and the assembly shall be inspected for dry spots, voids or other surface deformations.

3.2.20 Evaluation of Finished Surfaces

The cured Outer TF Leg Assemblies shall be checked for dimensional conformance with the drawings.

3.2.21  Installation of TF Coolant Tube Support:

The vendor shall install the TF coolant tube supports per details in PPPL approved drawings in Appendix I.

3.2.22  Final Inspection and Test:

The inspections and tests in section 4.0 paragraph 4.2.1 shall be performed once each Outer TF Coil Leg Assembly is in a deliverable state.

4.0 INSPECTION & TEST REQUIREMENTS

Mechanical and electrical inspections and tests shall be performed on each Outer TF Leg Assembly.  Results of the inspections and tests are to be provided to PPPL as part of the Process History (reference paragraph 5.12).  Seller shall notify PPPL a minimum of 5 days in advance to permit PPPL’s representatives to witness any of the inspections or tests, either in process or final acceptance.

4.1 RECEIVING INSPECTION-OUTER TF CONDUCTOR

4.1.1 Dimensional Measurements/Inspections

a. 	Conductors shall be checked to verify that length and cross sectional dimensions are within the required tolerance on the approved PPPL drawings.  They will also be inspected for excessive twist, warp and camber.

b. 	The dc resistance and temperature of each conductor shall be measured and recorded.

c.	A Rockwell hardness test shall be performed to check the conductor hardness.  These values shall become the baseline values for comparison checks throughout the fabrication process.

4.2 IN-PROCESS TESTING

4.2.1	Tests After Conductors have been Formed

a. 	Conductors shall be dimensional checked to verify that the formed conductors are within the required tolerance on the approved PPPL drawings.

b.	A Rockwell hardness test shall be performed on each formed conductor to verify the conductor hardness following the forming operation.

4.2.2	Inspection after Electrical Flag Installation

a.	A visual inspection of the finished joint shall be made to ascertain complete flow of Sil-Fos braze material into the joint area.  The joints shall be free from cracks and excessive porosity.

b.	A Rockwell hardness test shall be performed on each conductor end to verify the conductor hardness following the brazing operation.

4.2.3	Tests After Coolant Tube Installation

The soft-soldered edge coolant tube in each Outer TF Leg conductor must be tested for flow and pressure to verify the integrity of the coolant tube and to provide a base line for future tests.  These tests shall be performed using potable water.

a. 	Conductor Flow Tests:

A flow test shall be performed using calibrated pressure gauges.  Operational flow rates are as follows:

16.0	gpm		   1.7	psi pressure drop

b.	Hydrostat Tests:

A hydrostat test shall be performed using calibrated pressure gauges.  The pressure gauge shall be in increments no greater than 2 psi.

The test pressure shall be maintained at 600 psi without any change for at least ten minutes from the time the system was isolated from the pressure source, during which time there shall be no change in the pressure reading on the calibrated pressure gauge.

After the ten minute period, the pressure may be reduced to the 200 psi, and held for a sufficient time to permit inspection of the newly soldered coolant tube.  Record all test data.

4.2.4	 Tests After TF Conductor Assembly and Prior to Oven Cure

Once the Outer TF conductors have been assembled together as a Outer TF Coil assembly and ground wrapped, the insulation between each outer leg within the assembly shall be verified by applying a 500 Volt Megger for a minimum of 2 minutes between each individual turn and the remainder of the turns, which shall be shorted and grounded.  Insulation resistance shall be greater than 1000 MΩ.

4.3 FINAL ACCEPTANCE TESTS

These tests and inspections shall be performed once the TF Outer Coil Leg Assemblies are in a deliverable configuration.

4.3.1		Inspection:

A complete visual inspection and dimensional check shall be made according to applicable drawings.

4.3.2	Flow/Pressure Tests

A final flow and hydrostatic pressure test of the soft-soldered edge coolant conductor tubes shall be performed to verify that no tube damage has been done to the soft-soldered tube as a result of final processing of the Outer TF Coil Assemblies.  These tests shall be performed using potable water. Operational flow rates and pressure drops shall be within 10% of the data collected in section 4.2.2.

4.3.3	Final Electrical Tests

a. 	Coils shall be temporarily wrapped with conducting material (e.g. Al foil) to provide a ground plane.

b.  The insulation between each outer leg conductor within the assembly, and to ground, shall be verified by applying a 3 kV Megger between each individual conductor and the remainder of the conductors, which shall be shorted and grounded.  Insulation resistance shall be greater than 1000 MΩ.

c. 	The insulation between the coil assembly and ground shall be verified by applying a 3 kV DC Hipot between all turns, shorted together, and ground. Voltage shall be ramped up in 10 seconds, held for 60 seconds, and ramped down in 10 seconds.  Leakage current shall be recorded during ramping up and down, and once every 10 seconds during the 60 second interval.  Leakage current shall not exceed 10 µ Amps.

5.0 	QUALITY ASSURANCE REQUIREMENTS

5.1 	INSPECTION/ SURVEILLANCE/AUDIT BY PRINCETON

Authorized representatives of PPPL and the U. S. Government shall have the right at all reasonable times to visit the Subcontractor's premises and those of Subcontractor's suppliers during the performance of the procurement for the purposes of inspection, surveillance, audit and/or obtaining any required information as may be necessary to assure that items or services are being furnished in accordance with specified requirements. Such visits shall be coordinated with the Subcontractor's personnel to minimize interference with the normal operations of said premises. The Subcontractor shall make available records and documentation necessary for this function and shall provide all reasonable facilities and assistance for the safety and convenience of PPPL and/or U. S. Government representatives in the performance of their duties.   PPPL and the U. S. Government recognize the Subcontractor's right to withhold information concerning proprietary processes. The Subcontractor agrees to insert the paragraph above in each lower tier procurement issued hereunder.

5.2 	SUBCONTRACTOR’S RESPONSIBILITY FOR CONFORMANCE

Neither PPPL's review and/or approval of Subcontractor's documents nor PPPL's inspection of Subcontractor's items or services shall relieve the Subcontractor of responsibility for full compliance with requirements of the purchase order/contract.

5.3 	QUALITY ASSURANCE PLAN

Subcontractor shall submit  two (2) copies of its Quality Assurance Plan, describing the specific quality assurance and quality control procedures and practices to meet the requirements of this particular subcontract/purchase order, for PPPL review and approval. (Reference Section 6.2)

5.4 	MANUFACTURING, INSPECTION AND TEST PLAN (MIT)

Subcontractor shall prepare and submit for PPPL review and approval a Manufacturing/Inspection/Test Plan which identifies parts, sub-assemblies, etc.; shows their integrated flow into end items; and identifies critical manufacturing operations as well as inspections and tests. Preparing the Plan may include developing a flow chart and generating Process Sheets/Shop Travelers, etc. PPPL may designate selected manufacturing, inspection and/or test operations as mandatory “witness” points based on the MIT plan. Subcontractor shall provide PPPL with five (5) working days notice in advance of such witness points.

5.5 	INSPECTION AND TEST PROCEDURES

Inspections and tests shall be performed in accordance with written procedures referencing criteria for acceptance or rejection.  Each inspection and test shall be documented.

5.6 	DOCUMENT TRACEABILITY AND RECORDS

The Subcontractor shall maintain a system of documentation whereby objective evidence of required operations, inspections, examinations, and tests is systematically compiled, indexed and stored.  Such objective evidence may include "travelers"; and material test, certification, inspection, examination, test and discrepancy reports; which shall be complete, legible, and validated by responsible personnel and shall be traceable to subject items.

5.7 	EQUIPMENT/MATERIAL IDENTIFICATION AND STATUS

Material and equipment identification shall be maintained throughout the program and be traceable to the records. Status of acceptability shall be readily discernible through the Subcontractor's use of tags, stamps, serial numbers or other positive means.

5.8 	CALIBRATION OF TEST AND MEASURING EQUIPMENT

Inspections and tests shall be performed using properly calibrated measuring and test equipment. Subcontractor shall have in its possession the necessary equipment to perform the required inspections and tests. Calibration standards shall be traceable to the National Institute for Standards and Technology (NIST) or equivalent. Subcontractor shall impose these calibration requirements on subtier suppliers.

5.9      CONTROL OF SPECIAL PROCESSES

Subcontractor shall use trained and qualified personnel and qualified written procedures in accordance with specified requirements for the performance of certain special processes, including but not limited to, soldering, electronic assembly, brazing, welding, plating, heat treatment, nondestructive examination, etc. Copies of special process procedures shall be available for review by Princeton and submitted to Princeton for review and approval if requested.

5.10	NONCONFORMANCES

The subcontractor shall promptly identify and control nonconforming items or services.  Nonconforming items or services shall be positively identified, and segregated where possible, to prevent use.  The subcontractor shall document each nonconformance.  The written approval of Princeton is required prior to the use of the nonconforming item or service.

5.11 	RELEASE FOR SHIPMENT FORM

Subcontractor shall have a signed "Product Quality Certification and Shipping Release" Form to be provided by PPPL's Quality Assurance Representative prior to PPPL acceptance of procured items or services for full or partial shipment. PPPL reserves the right to refuse to accept shipments unless accompanied by a signed "Shipping Release" Form .

5.12 	PROCESS HISTORY

Subcontractor shall provide to PPPL a Process History which includes a compilation of documents, detailing the objective evidence of the acceptability of the work performed. The Process History shall  include as a minimum, but not be limited to, the following:

5.12.1 		Material Certifications

The Subcontractor shall submit copies of inspection reports, test data, and/or certifications from vendors, showing relevant chemical, mechanical and electrical properties of materials used, where applicable, as well as documents showing adherence to in-process requirements. Material certifications from sub-tier suppliers shall also be submitted.

5.12.2		 Inspection and Test Reports

Copies of the original reports of all required inspections, tests and examinations, properly validated by authorized personnel.

5.12.3		Fabrication Drawings

Copies of the final fabrication “as built”  drawings that include all approved changes made during the fabrication.

5.12.4		Certificate of Compliance (C of C)

Subcontractor's C of C, stating that the work performed conforms in every respect to the physical configuration and functional inspection/test requirements. The C of C shall be signed by Subcontractor's Quality Assurance (QA) Manager. Where Princeton-furnished material has been used by the Subcontractor, such certification shall also include the statement: "Material furnished by Princeton has been inspected by the Subcontractor and used by the Subcontractor as specified by Princeton with no unauthorized substitutions.

5.12.5	 	Nonconformance Reports

The subcontractor must submit copies of nonconformance or discrepancy reports, those affecting form, fit or function.

5.13 	PPPL RECEIVING/INSPECTION

PPPL will perform Receiving Inspection on items or services supplied by Subcontractor, using either a sampling plan or 100% inspection. Discrepant items or services will be rejected and returned to Subcontractor or reworked by PPPL.  Costs caused by rejects will be charged to Subcontractor.

6.0 DOCUMENTATION REQUIREMENTS

6.1 	PROPOSAL TECHNICAL DOCUMENTATION

a. 	Subcontractor shall submit as part of his proposal documentation in preliminary or outline form, the QA Plan (refer to section 5.3), the Manufacturing/Inspection/Test Plan (refer to section 5.4), and procedures for tests (refer to section 5.5).

b. Subcontractor shall submit as part of the proposal a preliminary description of the press mold, and assembly alignment fixtures which will be required to complete the assemblies.

6.2 	DOCUMENTATION AFTER CONTRACT AWARD AND PRIOR TO   	FABRICATION

a. 	Subcontractor shall submit (2 copies) final versions of the QA Plan (refer to section 5.3), the Manufacturing/Inspection/Test Plan (refer to section 5.4), and procedures for tests (refer to section 5.5) for PPPL review and approval within 30 days of contract award.

b. 	Subcontractor shall submit (2 copies) a final description of the press mold (refer to section 3.2.18)  and assembly alignment fixture (refer to section 3.2.15) for PPPL review and approval within 30 days of contract award.

c. 	Subcontractor shall submit (2 copies) of fabrication drawings for PPPL review and approval within 30 days of contract award.

6.3 	DOCUMENTATION PRIOR TO SHIPMENT

a. 	Subcontractor shall submit (2 copies) all documents identified in section 5.12, Process History.

b. 	Subcontractor shall submit an inventory of excess materials provided by PPPL.

7.0 	SHIPPING STORAGE AND HANDLING

a. 	The conductor coolant tubes shall be drained of all water and sealed for storage and subsequent shipment.

b. 	The Outer TF Leg Assemblies shall be individually crated for shipment.  The crate shall be built for moving on rollers, handling with slings from overhead cranes and forklifts.

c. 	The crate shall protect the Outer TF Leg Assemblies from shock, damage from load shift, and weather conditions, including precipitation.  Subcontractor’s name, shipper, purchase order number, Outer TF Leg Assembly (OTFL#) number (3.2.15) and gross weight shall be marked on the shipping container.  Subcontractor is responsible arranging shipment, and for the safe arrival of the PF coils at PPPL in Princeton , New Jersey, USA.

d. 	The Outer TF Leg Assemblies shall be prepared for shipment is such a manner as to ensure acceptance by common carrier for transportation at the lowest applicable rate and to afford protection from normal hazards of transportation.

8.0 	DELIVERABLES

The Subcontractor is responsible for delivering to PPPL the following:

a. 	All documentation listed in section 6.0.

b. 	Twelve (12) Outer TF Leg Assemblies which have successfully passed all acceptance tests and criteria.

c.	Balance of materials supplied by PPPL or which were procured by Subcontractor solely for this contract.

d. 	Fixtures or tooling fabricated solely for the fabrication of the Outer TF Leg Assemblies.

9.0	ATTACHMENTS:

APPENDIX I

Outer TF Coil Leg Assembly drawings

| Description | PPPL Drawing Number |
| --- | --- |

| Outer TF Leg Conductor Extrusion | CDC-1066 |
| --- | --- |

| Outer TF Leg Detail | EDC-1067 |
| --- | --- |

| Outer TF Leg Connector Details | EDC-1068 |
| --- | --- |

| Outer TF Leg Sub-assembly | EDC-1069 |
| --- | --- |

| Outer Leg Assembly | EDC-1070 |
| --- | --- |

| Outer TF Leg Coolant Tube Supt. Details | EDC-1082 |
| --- | --- |