National Spherical Torus Experiment

NSTX

Specification

for

Multipulse Thomson Scattering

Laser System

NSTX-SPEC-4X-58

Revision 3

9/03/98

PREPARED BY:		 Benoit LeBlanc________________________________

(Cog Engineer)

REVIEWED BY:	 William Slavin_________________________________

(Laser Safety Officer)

REVIEWED BY:	Frank Malinowski_______________________________

(QA)

APPROVED BY:	 David Johnson_________________________________

(WBS Manager)

PRINCETON UNIVERSITY

PLASMA PHYSICS LABORATORY

(PPPL)

PO BOX 451

PRINCETON, NJ 08543

609-243-2000

| Revision | Date | TRB | Description of Change |
| --- | --- | --- | --- |

| 0 | 7/29/98 |  | Initial Document, John Robinson |
| --- | --- | --- | --- |

| 1 | 8/10/98 |  | Section 2.0, 3.1, 4.0, B. LeBlanc |
| --- | --- | --- | --- |

| 2 | 8/81/98 |  | Section 2.0, 3.1, 4.0, B. LeBlanc, 30 Hz only |
| --- | --- | --- | --- |

| 3 | 9/3/98 |  | Section 5.0, BPL as per Malinowski |
| --- | --- | --- | --- |

1.0	SCOPE

Subcontractor shall supply a standard commercial Nd:YAG laser system(s) for use in the Multipulse Thomson Scattering System (MPTS) diagnostic on the National Spherical Torus Experiment (NSTX). All specifications shall be adhered to unless there is prior written agreement to the exceptions.

2.0 APPLICABLE DOCUMENTS

NONE

3.0 REQUIREMENTS

3.1 PERFORMANCE REQUIREMENTS

3.1.0 LASER MODE OF OPERATION AND USAGE INFORMATION

The laser will be used as part of a Thomson scattering diagnostic system, where laser light scattered off the plasma electrons is collected to measure the electron temperature and density. The laser will be operated in a burst mode. There will be typically 50 bursts daily, each separated by about 10 minutes. Each burst will consist of a 125-second long, fixed-frequency sequence of pulses. Only the last 5 seconds of the sequence will be used for measurement and all performance characteristics listed below apply to this last 5-second interval, which we shall call the run sequence.

3.1.1 PERFORMANCE CHARACTERISTICS FOR RUN SEQUENCE

Parameter	Value	Installation Requirement (see 4.0)

Wavelength:	1064.2 nm

Wavelength Stability:	≤0.1 nm

Linewidth:	≤1.0 cm-1

Repetition Rate1	30 Hz	√

Energy Output2	≥1.6 J	√

Output beam diameter	1 cm	√

Pulse Width3:	8 ns ≤ pulse width ≤ 15 ns	√

Energy Stability4	≤ 5% of mean value	√

Divergence5:	≤ 0.5 mradian	√

Beam Pointing Stability6:	< 50 μradians  	√

Q-switch Timing Jitter7:	≤ 3 ns	√

Polarization8:	Linear vertical	√

Computer Interface9:	IEEE-488, RS-232	√

1	The laser shall operate at a constant frequency and both flashlamp and Q-switch shall be triggerable by internal or external clock, selected by a switch setting.

2	Average reading per pulse into a calibrated energy meter.

3	Full width at half maximum

4	Individual pulse energy shall deviate by ≤5% from the mean energy value during the 5 sec. run time.

5	Full angle defined by 86% of full energy, from a 1-cm diameter collimated beam, passing through a circular aperture of diameter d located at the focal point of a converging lens of focal length f, such that d = f x divergence, and f ≥ 5 m.

6	The beam center shall deviate from mean by less than 50 μradians over 10 hours of burst operation as described in 3.1.0.

7 	RMS jitter in timing between Q-Switching and laser output.

8	More than 95% of energy in designated polarization state.

9	ASCII command structure (as in IEEE-488) is to be preferred.  LabVIEW drivers requested.

3.1.2 ENVIRONMENTAL

To be contained in a temperature controlled environment at <±3oC.

3.1.3 MAINTAINABILITY

Ease of service is an important consideration. The successful vendor must demonstrate its ease of removal for standard user serviceable parts (flashlamps, filters, rods, optics, etc.).

3.1.4 SAFETY

Shall comply with Federal Regulations (21 CFR Subchapter J) as to laser safety and health. Must incorporate external interlock safety access. Must have a final beam shutter capable of handling the optical output of the device.

3.2 EQUIPMENT DEFINITION

3.2.1 GENERAL DESIGN FEATURES

Shall be no larger than:

Head:			48”L  X 21”W  X 13”H

Power Supply:		31”L  X 25”W   X 37”H

3.2.2 ELECTROMAGNETIC INTERFERENCE AND SUSCEPTIBILITY

Product shall be subject to low EM fields (10s of gauss max.) and must be capable of working as specified. Product will not emit significant amounts of EFI/RFI and with covers closed must meet government standards in those areas.

3.2.3 IDENTIFICATION AND MARKING

Must include standard class IV laser identification, std. OSHA, NEC & NESC labels on the equipment.

4.0 TEST REQUIREMENTS

Upon installation vendor will perform to PPPL’s satisfaction installation tests verifying specifications indicated in Section 3.1.1 by symbol “√”. This verification will occur before acceptance is complete and before final payment is authorized to vendor.

5.0 QUALITY ASSURANCE REQUIREMENTS

5.1 INSPECTION/ SURVEILLANCE/AUDIT BY PRINCETON

Authorized representatives of PPPL shall have the right at all reasonable times to visit the Subcontractor's premises and those of Subcontractor's suppliers during the performance of the procurement for the purposes of inspection, surveillance, audit and/or obtaining any required information as may be necessary to assure that items or services are being furnished in accordance with specified requirements. Such visits shall be coordinated with the Subcontractor's personnel to minimize interference with the normal operations of said premises.  The Subcontractor shall make available records and documentation necessary for this function and shall provide all reasonable facilities and assistance for the safety and convenience of PPPL representatives in the performance of their duties.  PPPL recognize the Subcontractor's right to withhold information concerning proprietary processes. The Subcontractor agrees to insert the paragraph above in each lower tier procurement issued hereunder.

5.2 SUBCONTRACTOR'S RESPONSIBILITY FOR CONFORMANCE

Neither PPPL's review and/or approval of Subcontractor's documents nor PPPL's inspection of Subcontractor's items or services shall relieve the Subcontractor of responsibility for full compliance with requirements of the purchase order/contract.

5.3 CHANGES TO PRINCETON APPROVED DOCUMENTS

Revisions or changes by the Subcontractor to documents approved by PPPL shall be reviewed and approved by PPPL prior to use.

5.4 SUBCONTRACTOR QUALITY ASSURANCE PROGRAM

The Subcontractor shall establish and maintain an effective Quality Assurance Program and the general criteria to be satisfied by the Program. The objectives of such a program are to assure that the Subcontractor's work meets the required level of quality and is performed in accordance with contractual requirements.

Subcontractor's quality assurance function shall have sufficient authority and independence to identify quality problems, verify conformance of supplied items or services to specified requirements and obtain satisfactory resolution of conflicts involving quality.

5.5 ACCEPTANCE INSPECTIONS AND TESTS

Each item or service to be delivered to PPPL shall be inspected and tested by the Subcontractor to verify that such items or services meet PPPL's requirements.  Qualified personnel who have not performed or directly supervised the actual work shall perform inspections and tests in accordance with written procedures referencing criteria for acceptance or rejection.  Adequate records shall be maintained and available for PPPL's review.

5.6 ACCEPTANCE TEST PROCEDURES

The Acceptance Test Procedures (ATPs), including pass/fail criteria, required to demonstrate conformance to PPPL's requirements after installation shall be prepared by Subcontractor and two (2) copies submitted to PPPL for review and approval with the proposal.

5.7 CALIBRATION OF TEST AND MEASURING EQUIPMENT

Inspections and tests shall be performed using properly calibrated measuring and test equipment. Subcontractor shall have in its possession the necessary equipment to perform the required inspections and tests. Calibration standards shall be traceable to the National Institute for Standards and Technology (NIST) or equivalent. Reference standards shall only be used for calibration and verification unless it is demonstrated that their performance as reference standards has not been invalidated. Subcontractor shall impose these calibration requirements on subtier suppliers.

5.8 ENVIRONMENTAL, SAFETY, AND HEALTH, REQUIREMENTS

The Quality Assurance Program shall require that the Subcontractor's design, fabrication, installation and checkout activities shall take into account environmental, safety, and health requirements with respect to personnel and radiological hazards, and generation and disposal of hazardous waste.

5.9 RELEASE FOR SHIPMENT FORM

Subcontractor shall have a signed "Product Quality Certification and Shipping Release" Form to be provided by PPPL's Quality Assurance Representative prior to acceptance of procured items or services for full or partial shipment. PPPL reserves the right to refuse to accept shipments unless accompanied by a signed "Shipping Release" Form.

5.10 PROCESS HISTORY

Subcontractor shall provide, at or prior to delivery, PPPL with two (2) copies of the Process History, a compilation of documents, detailing the objective evidence of the acceptability of the work performed. The Process History shall include as a minimum, but not be limited to, the following:

5.10.1 AS-BUILT CONFIGURATION

The Subcontractor must submit drawings and a listing of the as-built configuration of each delivered item.  The listing may be defined by drawing and revision numbers, unique parts lists or other means of positive identification.

5.10.2 INSPECTION & TEST REPORTS

Copies of the original reports of all required inspections, examinations, and tests, properly validated by authorized personnel.  One copy to be provided with the Shipping Release request.

5.11 GENERAL PURCHASE ORDER REQUIREMENTS

Material(s) and/or product(s), including those components, parts, and materials, which are permanently installed into systems, sub-systems, and/or assemblies, etc. furnished under this purchase order/subcontract, shall be new and unused. Parts and components that have been rebuilt, refurbished, or modified are specifically prohibited unless approved by PPPL in writing.

Subcontractors (Suppliers) found to be supplying any of the following will be reported to the Office of the Inspector General, US Department of Energy.

•	Remanufactured, rebuilt, or used parts represented as new

•	Counterfeit parts (fraudulently labeled or marked with another manufacturer’s name).

•	Misrepresented parts.

6.0 HANDLING, SHIPPING AND STORAGE

Packaging shall be sufficient to assure that materials and items are adequately protected from damage or deterioration in shipment and during environmentally controlled warehouse storage. Laser shall be shipped in the normal safest mode (usually air-cushioned and environmentally controlled truck). Labeling shall be specific as to handling and anti-tamper labels are preferred for opening, tipping, rough handling, etc.  Inspection in the warehouse upon arrival is mandatory by laser personnel (preferably before the trucking firm leaves).

7.0 WARRANTY

Standard warranty is acceptable and should be at least 2 years on equipment and a shorter time (not to be less than 90 days) for optics, crystals and coatings (should be from the time of installation and acceptance).

8.0 DELIVERABLES

8.1	Test procedures-with the proposal

8.2	Product Quality Certification and Shipping Release Form, top portion completed-prior to shipment

8.3	Manufacturing Test report(s)-with the Release for Shipment request

8.4	Laser system-as per proposal

8.5	Instruction books-with the laser system

8.6	Acceptance Test Reports-on successful completion of post-installation tests

ATTACHMENTS

"Product Quality Certification and Shipping Release" Form

PRINCETON UNIVERSITY

PLASMA PHYSIC LABORATORY—PPPL

| PRODUCT QUALITY CERTIFICATION AND SHIPPING RELEASE |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

| PROJECT | ITEM DESCRIPTION |  |  | SHIPMENT NUMBER |  |
| --- | --- | --- | --- | --- | --- |

| PPPL SUBCONTRACT/ ORDER NO. | REV. | ITEM NO. | SUPPLIER REFERENCE NO. | REV. | QUANTITY SHIPPED |
| --- | --- | --- | --- | --- | --- |

| SUPPLIER'S CERTIFICATION This is to certify that the products and services identified herein have been produced under a controlled quality assurance program and are in conformance with the procurement requirements including applicable codes, standards and specifications as identified in the above-referenced documents unless noted below. Any supporting documentation will be retained in accordance with the procurement requirements. SIGNED: ______________________________________ DATE: __________________________ TITLE: _______________________________ COMPANY: _______________________________ |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

| PPPL (AUTHORIZED REPRESENTATIVE) SHIPPING RELEASE This is to certify that evidence supporting the above Supplier's Certification statement has been audited and no product/service nonconformances from procurement requirements have been found unless noted below. This product/service is hereby released for shipment. This section serves as the Quality Assurance release for the above-described product for shipment. It does not constitute an acceptance thereof and does not relieve the Vendor, Manufacturer or Contractor of any and all responsibility or obligation imposed by the purchase contract. It does not waive any rights the Purchaser may have under the purchase contract, including the Purchaser's right to reject the above described material upon discovery of any deviations from requirements of the purchase contract, drawings and specifications. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

| NONCONFORMANCES FROM PROCUREMENT QUALITY REQUIREMENTS: |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

| REMARKS/PRODUCT SERIAL NUMBERS: |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

| BY PPPL QA REPRESENTATIVE (OR DESIGNEE) |  |  |  | DATE |  |
| --- | --- | --- | --- | --- | --- |