4X-970828-RK-01

TO: DISTRIBUTION

FROM: R  KAITA
SUBJECT:  EVALUATION OF CURRENT PROFILE DIAGNOSTICS

Reference: Action Item # 33

This memo brings to closure Action Item #33 from the from the Engineering Cost and Schedule Review of July 30 - August 1, 1996. The concern is stated as follows.

“Because of the low field strengths in NSTX the MSE diagnostic may not be effective.”

To address this, a Current Profile Diagnostics Working Group Meeting was held on July 16, 1997 to evaluate improvements to MSE and alternative techniques for current profile measurements. Recommendations are provided in the summary of the meeting (Appendix A). The agenda (Appendix B) and list of attendees (Appendix C) are also appended to this memorandum.

cc:

D Johnson	P Heitzenroeder 	S Kaye	C Neumeyer

M Ono	M Peng		C Such

NSTX File

Appendix A

NSTX Current Profile Study Group Meeting

Princeton Plasma Physics Laboratory

July 16, 1997

David Johnson, Robert Kaita  -  PPPL

Introduction

Capability for measuring j(r) has become extremely valuable in several areas of tokamak physics, and it is expected to be even more important in ST research.  However, the ST environment imposes new challenges on such measurements.  Because the toroidal field of the ST is reduced by up to an order of magnitude, the motional Stark splitting of the collisionally excited beam emission is very small and the different polarization components can be smeared to the point that conventional MSE polarimetry becomes ineffective.  This issue was discussed at the NSTX Research Forum, held February 5 - 7, 1997 at PPPL, which recommended that a Study Group be formed to define development needed to measure j(r) on NSTX.  In April 1997, an announcement broadly circulated to experts in the fusion community invited participation in a one-day Study Group Meeting to recommend the j(r) diagnostic development needed for NSTX.   A list of attendees and the agenda for the meeting are attached.

This document summarizes the results of the meeting from the viewpoint of NSTX diagnostics.  It highlights all the presentations and discussions without attempting to be complete.  A complete set of presentation transparencies will be sent to participants and is available to others on request.  Because the j(r) diagnostic is not within the approved baseline project budget, and because the resources available for diagnostic upgrades are still being defined, our recommendations cannot be as specific as we would like.  However, it is encouraging to point out that the time scale required to design, implement, and debug many of the systems discussed is 2-3 years if work on them begins immediately.  Significant R&D appear warranted for certain potentially attractive approaches, requiring funding for their development.

Summary and Recommendations

Equilibrium Reconstruction

Presentation  	R. Kaita, PPPL

As an introduction to the meeting, Bob Kaita reminded the group that j(r) diagnostics do not actually measure j(r) but rather provide constraints to an equilibrium calculation used to determine j(r).  Echoed frequently in many discussions was the point that the optimum design and the ultimate utility of

a j(r) diagnostic can be judged by the sensitivity of the equilibrium reconstruction to the diagnostic constraints.  Another useful comment from Ray Fonck was that participation of those concerned with MHD theory on NSTX should be encouraged in these judgments.

Recommendation:  The project should develop the capability (for example, using TSC and EFIT) to study self-consistently the sensitivity of an equilibrium reconstruction to various types of diagnostic input data, including both local [such as MSE data on the pitch angle γ(R) ] and line integral (such as Faraday rotation data on ∫nB⋅dl ) measurements, for a variety of expected NSTX equilibria.  Development of these tools will naturally enhance data analysis capability.

Faraday Rotation

Presentations  	T. Peebles - UCLA

C. H. Ma - ORNL

D. L. Brower - UCLA

H. K. Park - PPPL

There were several good presentations on Faraday rotation measurements, which both recalled system design and performance issues with past systems on C-MOD, TFTR, and TEXT, and also considered possible implementation on NSTX.  These multichord systems are envisioned to provide both interferometric measurements of ∫ndl, and polarimetric measurements of ∫nB⋅dl. Considering the NSTX parameters, the requirements for adequate phase precision in these two measurements and, finally, a desire to minimize misalignment due to laser beam refraction at high densities, the optimum laser wavelength is in the range 100-200 μm.

We were reminded that the Faraday rotation angle is a chord-integrated measurement, which requires two differentials to obtain the current profile j(r).   This seriously limits the accuracy for detecting small changes in j(r), compared to the MSE technique.

A number of geometries for Faraday rotation were explored.  The traditional vertical view, such as that used on TEXTOR, TFTR, TEXT and other devices, is not available on NSTX without large changes to the vacuum vessel.  Assuming that the vertical access is made available, vertical chords from the magnetic axis out to 2/3 of the minor radius at the midplane could become possible, still leaving the views for the outer 1/3 unavailable.

An interesting possibility for NSTX is a tangential fan of interferometer/polarimeter chords.  This fan, in order to include the core region, would be on the midplane, although partial profiles could be obtained with fans off the midplane.  For easy inversion and for cancellation of the poloidal magnetic field component, the fan should be in a horizontal plane.

Access requirements for such a system are relatively modest.  Beams would enter on a single midplane port and retro-reflectors would be used to direct the return beams out the same port.  Such a system would measure ∫neBφdl  and ∫nedl. With a sufficient number of chords, radial profiles of ne and Bφ can be obtained by spatial inversion.  This dual use capability is attractive and further enhanced by the high time resolution, which will be very useful on NSTX, as they have been in many tokamak experiments.   Measurements of Bφ(R) will also be very important on NSTX, due to the expectation of large paramagnetic and diamagnetic modifications of the vacuum toroidal field, and will provide additional constraints on equilibrium reconstruction.

Recommendation:  Due to the importance of knowing Bφ(r) on NSTX, study the feasibility and cost of a tangential multichord FIR interferometer/polarimeter system for NSTX to provide time resolved radial profiles of ne and Bφ.

Flux Surface Mapping with X-Ray Pinhole Camera

Presentation 	B. Stratton - PPPL

In an axisymmetric discharge the flux surface shapes as indicated by soft x-ray emission can be used to constrain the equilibrium analysis and derive j(r), assuming ne and Te are constant on a flux surface.  Tangential 2-D soft x-ray imaging was used on PBX-M to demonstrate this technique.  The hardware needed is relatively straightforward and cheap, given that significant tangential access is planned for NSTX.  Most of the effort here will be in the data analysis and integration with the equilibrium reconstruction.  There were concerns voiced about the uncertainties introduced by large run-away electrons, plasma rotation, or MHD activity.   A prototype system is planned for the Pegasus ST, which should provide very valuable information for this diagnostic for NSTX.

Recommendation:  Encourage the diagnostic development for soft x-ray imaging of flux surfaces on Pegasus and plan allocation of port access for a similar system on NSTX for implementation guided by the Pegasus experience.

Motional Stark Effect

Presentations	N. Bretz - PPPL

F. Levinton - FP&T

P. Carolan - Culham

The MSE technique recently revolutionized tokamak experimentation by providing local, time-resolved measurements of the pitch angle

γp=tan-1(Bp/BT).  These data give direct indications of the magnetic axis location and q(0), and provide sensitive constraints for determining j(r) and q(r) in the equilibrium reconstruction.   Such information has resulted in significant physics insight as well as improvements in tokamak performance.

At several Tesla in magnetic field, the vBxBT Stark splitting of the Hα multiplet is spectrally larger than various Doppler line broadening effects.  The splitting is also large enough that standard multicavity interference filters can be used to isolate the polarization components for effective polarimetric determination of the pitch angle.  However, at the NSTX toroidal field as low as 0.3T, the Stark splitting is smaller than normal Doppler broadening and also smaller than the passband of typical interference filters.  Thus, to extend this powerful technique to the ST environment, both the Doppler broadening effects and the effective spectral passband of the analyzer must be reduced, in each case by a factor of at least 3-4.  Norton Bretz commented on these issues in his presentation. Paddy Carolan described methods for calculating these effects on various devices in his talk on the spectral and polarization smearing that can affect MSE measurements.

Fred Levinton offered a number of innovative ideas at the meeting to solve these problems.  He proposed enhancements in the traditional approach, which uses collisionally induced fluorescence (CIF), and also proposed a more radical departure using laser induced fluorescence (LIF), which has a number of potential advantages.  For NSTX, he estimates that enhancements to the CIF method would be adequate for determination of γp  and⎟B⎟ = (BT2 + BP2)0.5.  The LIF MSE technique looks very attractive as a diagnostic for low field devices in general, and can be effectively used with a small diagnostic neutral beam (DNB) in the absence of a heating beam.  In fact, installing a DNB/LIF MSE system on NSTX in combination with a heating beam based, CIF MSE system would provide enough information to determine Er(r), a very important addition to the confinement measurement of essentially all toroidal plasmas.

In the NSTX situation, the dominant Doppler line broadening mechanism is the angular spread due to the finite viewing aperture.  Fred Levinton pointed out that this spread is primarily along one dimension in the aperture plane (horizontal in the case of NSTX). He suggested that the broadening can be reduced by maintaining the same aperture area while shrinking the aperture in the horizontal dimension and expanding it in the vertical dimension.  This will require a modification to the conventional input optics design, but should then enable use of a TFTR beamline for MSE.

Ray Fonck suggested that other well known techniques in spectroscopy of beams be investigated, which compensate for the angular spread in the viewing optics.

To address the need for a narrower spectral passband, both Fred Levinton and John Howard, of the Australian National University offered ideas using bi-refringent spectral filters, based on a Lyot design currently utilized in solar spectroscopy and dye laser tuning.  New development

appears necessary in this area, with the goal of an effective passband of 0.050 - 0.075 nm.  John Howard suggests doing the spectral analysis in a device he is developing called a 'modulated solid state interferometer' (MOSS), which measures the temporal coherence (inverse spectral bandwidth) of the incoming light.  It is based on electro-optically modulating a bi-refringent crystal in a Lyot configuration.

Fred Levinton has modeled the performance of a system, with the new aperture shape and a 0.075 nm filter passband, and obtains polarization fractions of Pf=0.2-0.4.  To maintain the precision of the MSE system on TFTR, which has a Pf~0.8, one must increase signal strength on the NSTX system by ~ x10 to compensate for the degraded Pf.   He proposes increasing the optical throughput by using more fibers, faster optics, and larger, more efficient detectors.

Fred Levinton described how laser induced fluorescence (LIF) can lead to further improvements.  He proposes to use a tunable CW ring dye laser coaxial with either a heating beam or a DNB.  With this geometry, the Doppler line broadening is minimized to ~ 1 mÅ, and a saturation power of only ~ 1 watt is needed in a 6 mm dia. laser beam.  The linear polarization of the beam would be rotated to measure the pitch angle, eliminating the need for photoelastic modulators (PEMs) in the light collection system.  He also proposed scanning the laser wavelength (or alternately the beam voltage) to measure⎟B⎟ to a precision of 0.1%, providing the potential for accurate determinations of total plasma pressure.  This scheme is well matched to a very compact DNB, with a ~ 1 cm beam diameter.

Over the last several years, the MSE technique has evolved to become the method of choice in tokamak research.  Because it is a local measurement only one integral removed from j(r), MSE data provides very powerful constraints on the equilibrium reconstruction.  We are excited about the prospects for extending this technique to the low field ST situation.

Recommendation:  In support of a CIF MSE system using the NSTX heating beam:

a) Develop wide field, narrow passband bi-refringent filter technology probably based on Lyot design.  Also consider modulated bi-refringent system for spectral analysis

b) Develop the design of high throughput collection optics with viewing aperture optimized to reduce Doppler line broadening

c) Develop design of high throughput analyzer optics and evaluate large area APD detectors.

It was seen that the enhanced CIF MSE system at present is the most promising development path for an effective j(r) diagnostic for NSTX, particularly assuming the possible installation of an upgrade NBI system.

At lower priority, we support the development of LIF MSE technology in conjunction with a DNB, to be considered as a possible future upgrade diagnostic for NSTX.  It holds promise to give higher spatial resolution for γp measurements and more precision in⎟B⎟ measurements.  In combination with a CIF MSE system viewing the heating beam, it could provide measurements of Er.  An increased priority should be applied to the LIF MSE development in the event that the heating beam for NSTX is delayed, or a technical impediment arises in the development of the CIF MSE system.

Heavy Ion Beam Probe

Presentations	T. Crowley - RPI

F. Jobes - PPPL

There were two proposals to use the trajectory of a heavy ion beam to deduce the internal magnetic field.  Tom Crowley proposed injecting a 250 keV Na+ beam through a dome port and observing the secondary beam emerging through a midplane port.  The most straightforward scheme would be to scan the primary beam in a radial plane, and observe the secondary with a toroidally extended detector at a fixed poloidal position within the TF coils on the midplane. The detector would be sensitive to the secondary beam intensity and to the toroidal component of the beam velocity.  This would allow determination of the poloidal flux Ψ along a trajectory of ionization sites distributed through the plasma.  Also, fluctuations in the beam intensity provide a measure of the density fluctuations at these sites.  Signal strength is not a dominant issue in these measurements, and excellent time response (~ 10 μsec) should be possible.  Analysis would involve an iterative process of calculation of trajectories based on an equilibrium, and then adjusting the equilibrium to fit the experimental data.  A HIBP system using a  500 keV Tl was used successfully on TEXT to infer j(r).  Highly noncircular flux surfaces and the expectation that BP ~ BT at high β would both complicate this analysis on NSTX.   Further study is needed to evaluate the stiffness of constraints on 'typical' NSTX equilibria provided by such data.   R&D is also needed to develop suitable velocity detectors.

Buzz Jobes proposed to use the visible light emitted by the primary beam to accurately determine its trajectory in 3D as a probe of BP and BT.  He would use a 120 keV Na+ beam injected radially and slightly upward to probe through the magnetic axis.   The beam would be viewed at right angles from above and from the side with high resolution CCD cameras.   Again an iterative procedure involving a comparison of calculated and measured trajectories and subsequent adjustment of the equilibrium would be used.  More work is needed here to estimate the brightness of the emission, the expected uncertainties in the trajectory measurement, and the stiffness of the resulting constraints on various equilibria.

In the early stages of NSTX operation, limited resources are likely to prevent the installation of a large HIPB facility like that used on TEXT.  However, because of the wealth of information that can be obtained with such a facility, it should be considered as a possibility in the future.  The availability of more compact Na+ beams of appropriate energy should also be examined.

Recommendation:  We encourage more work to define the stiffness of constraints resulting from heavy ion beam trajectory data from both of the above techniques for various NSTX equilibria.

Other Techniques

Through the course of discussion and through submissions from researchers unable personally to attend the meeting, a number of other techniques were briefly mentioned.

Thomson scattering - Measurement of the drift velocity by looking at the spectral shift in the thermal distribution from Thomson scattering has been done on several devices, most recently with some success on RTP.  However, the time resolution and precision of these measurements have typically been inadequate to compete with other techniques.

Transient internal probe - (see submission by T. Jarboe - Univ. Washington) This technique involves using a magneto-optic probe accelerated at high speed through the plasma with a gas gun.  The local field is then determined from the Faraday rotation of a laser beam reflected from the probe as it traverses the plasma.  This technique was successfully demonstrated on the HIT device, and was suggested for use during low temperature startup phase of NSTX.  The transient, perturbing nature of this technique limits its usefulness for NSTX.

Fizeau interferometer - (see submission by A. Mahdavi - General Atomics) This is a new idea proposed as a diagnostic to supplement MSE on DIII-D for core j(r) measurements, particularly at high density, where poor beam penetration makes MSE difficult.  The technique involves measuring the phase shift difference in two counter-propagating FIR laser beams following identical paths on a midplane toroidal chord.  Since the beams are strictly counter-propagating, the normal phase shift cancels out, if the medium is stationary.  There is a relative phase shift proportional to ∫j(r)dl, which can be spatially inverted to give j(r).

Recommendation:  Study the physics, feasibility, and cost of a tangential multichord FIR Fizeau interferometer system for NSTX to provide time resolved measurements of ∫j(r)dl.

Correlation reflectometry arrays - (see submission by T. Peebles - UCLA) Relying on the fact that the correlation length for fluctuations is much larger in the parallel than the perpendicular directions, this technique would use the correlation of fluctuation data from two poloidal arrays of reflectometers located on the same midplane port to determine the pitch angle. Obtaining profiles would require sweeping the reflectometer frequency during the shot to probe different radial positions.

O- and X-mode edge reflectometry - (see submission by G. Hanson - ORNL )   An edge reflectometer for density measurements is planned as part of the RF program on NSTX.  It may be possible to configure this system to launch both O- and X-mode waves to probe the edge region, which, in principal, could yield both ne(r) and⎟B⎟.

Passive Zeeman spectroscopy/polarimetry - (see submission by J. Howard - Australian National Univ.)  Variants of this technique have been used on TEXT, TFTR, and JIPP T-IIU to obtain relatively crude measurements of the internal magnetic fields.  These are line-integrated measurements weighted by the line intensity distributions.  Seeding with impurities has generally been necessary to increase brightness.  It is not clear whether these techniques are feasible at the low fields in NSTX.

Appendix B

CURRENT PROFILE DIAGNOSTIC STUDY GROUP MEETING AGENDA

July 16, 1997

Welcome and General Remarks

8:30	R. Kaita		PPPL

General Remarks on Equilibrium Reconstruction

Refractive Index and X-Ray Measurements

8:45	1) W.A. Peebles	UCLA

Toroidal Magnetic Field Measurements on NSTX and their Relevance to Current Profile Determination

9:15	2) C. H. Ma		ORNL

Design Considerations of an Interferometer/Polarimeter for  Measurement of Current Profile in NSTX

9:45	3) D. L. Brower	UCLA

FIR Polarimetry on NSTX

10:15 - BREAK

10:30	4) H. K. Park	PPPL

Poloidal and Toroidal Field Measurements with Faraday Rotation Methods

11:00	5) B. Stratton	PPPL

Soft X-Ray Imaging for Equilibrium Reconstructions

11:30 - DISCUSSION OF REFRACTIVE INDEX AND X-RAY MEASUREMENTS

12:00 - LUNCH

Spectroscopic Techniques

13:30	7) N. Bretz	PPPL

MSE Measurement of Current Profile

14:00	8) F. M. Levinton	Fusion Physics and Technology, Inc.

A Low Magnetic Field, Motional Stark Effect(MSE) Diagnostic, Providing Magnetic Field Pitch Angle, Radial Electric Field, and Kinetic Pressure

14:30	9) P. Carolan	Culham Laboratory

Treatment of Spectral and Polarisation Smearing in a Motional Stark Effect Diagnostic

15:00 - DISCUSSION OF SPECTROSCOPIC TECHNIQUES

15:30 - BREAK

Beam Probe Measurements and Other Techniques

15:45	10) T. P. Crowley		Rensselaer Polytechnic Institute

Application of Heavy Ion Beam Probe Magnetic Field Measurement

Techniques to NSTX

16:15	11) F. Jobes			PPPL

Heavy Ion Beam Probe Diagnostic for NSTX

16:45	12) J. Howard (poster  - highlights summarized by R. Kaita)	ANU

Laser Polarimetry and Zeeman Spectroscopy

17:00 - DISCUSSION OF BEAM PROBE MEASUREMENTS AND OTHER TECHNIQUES

17:30 - ADJOURN

Appendix C

List of Attendees

| R. Kaita - PPPL | P. Carolan - Culham Lab. | S. Medley - PPPL | K. Thomassen - LLNL | B. Stratton - PPPL |
| --- | --- | --- | --- | --- |

| C. H. Ma - ORNL | T. P. Crowley - RPI | F. Paoletti - Col. Univ. | T. Carter - PPPL | N. Bretz - PPPL |
| --- | --- | --- | --- | --- |

| W.A. Peebles - UCLA | F. Jobes - PPPL | M. Peng - ORNL | J. B. Wilgen - ORNL | D. Markevich - DOE |
| --- | --- | --- | --- | --- |

| D. L. Brower - UCLA | D. Johnson - PPPL | M. Ono - PPPL | G. R. Hanson - ORNL | W. Dove - DOE |
| --- | --- | --- | --- | --- |

| H. K. Park - PPPL | K. Hill - PPPL | R. Fonck - Univ. of Wis. | C. Forest - Univ. of Wis. | F. M. Levinton - FP&T |
| --- | --- | --- | --- | --- |