### 908 Digitizer Record Fields

Rev 0   	11NOV98

P.  Sichta

# Introduction

This document will describe the fields of the new EPICS record for BiRa 908 or Aeon 3232 transient digitizers. The 3232 in 908-compatability mode is supported at this time.

908 EPICS Record Fields

The record will not need any device support.  Its function is to hold information that can be used by an acquisition program or plotting program (probably on the host).  There will be  PV Links so that some of its fields can be loaded from other EPICS records.

| Field Description | Field ID | Field Type | PP | GDCT Menu |
| --- | --- | --- | --- | --- |
| Description … and other “standard” record fields | --- | --- | --- | --- |
| This field is for a database link to the VAL field of a record which reads CAMAC, F0.A0. This is the status register of the 908. The link should be NPP since the CAMAC record will scan periodically. | LMST | IN_LINK | N | Y |
| This field is for a database link to the VAL field of a record which reads CAMAC, F0.A2. It will provide the number of valid samples. | LMNV | IN_LINK | N | Y |
| This field should connect to a source which will provide the time of trigger, relative to T(0). This is typically linked to the VAL field of a 404 Timing Module “SUB” record. | LTOT | IN_LINK | N | Y |
| This field is for a database link to the VAL field of a record which reads CAMAC, F0.A1. It will provide the number of post-trigger blocks (block=16 samples) which were taken for in pre-trigger mode. This field is optional and is required only for 908’s in PreTrigger Mode. Note that the Pre/Post Trigger status in MSTS disappears after an unload/readback, so post trigger samples must have been setup to be at least 16 (samples, 1 block) to inform the software that the data was taken in PreTrig mode. This field is zero for post-trigger acquisition. | LPTB | IN_LINK | N | Y |

| Same as LMST, this field is optional an is required only when the internal sample clock is not used. The record supplying the value should provide the sample period in seconds. | LEXT | IN_LINK | N | Y |
| --- | --- | --- | --- | --- |
| Value of the Master 908 Status word derived from LMST. | MSTS | LONG | N | N |
| Value of the 908 Operating State derived from LMST | OPST | STRING | N | N |
| Value of the 908 Operating Mode derived from LMST | OPMD | STRING | N | N |
| Value of the 908 Input Voltage Range derived from LMST. Note that the 3232 digitizer has no +/- 2.5 volt range. If the indication is +/- 2.5 for an Aeon 3232 then that means it is actually a +/- 10 volt input range. | INPR | STRING | N | N |
| Value of the Sampling Period derived from LMST, or if external clock mode from EXCK. | SPRD | DOUBLE | N | N |
| Inverse of SPRD. | SRAT | STRING | N | N |
| Value of external sample period (in seconds) from LEXT. | EXCK | DOUBLE | N | N |
| Value of the Number of Active Channels derived from LMST. | NCHN | SHORT | N | N |
| Total Digitizer Memory derived from LMST | MEMR | LONG | N | N |
| Value of the Master 908 Num. Valid Samples derived from LMNV. | MNVS | LONG | N | N |
| Value of the Time of Trigger derived from LTOT. | TOT | DOUBLE | N | N |
| Value of the Post Trigger Samples derived from LPTB, multiplied by 16 (16 samples per block). | PTS | LONG | N | N |
| Read-In-Progress flag. This will be initialized to zero. When a data channel is to be read, this should first be checked and set to one by the readin-program prior to the F16.A1 unload request. After all data has been read, the program must clear the flag. This field is meant to keep the data read process “atomic”. | RIP | SHORT | Y | N |

| This is a specially-formatted string to inform plotting and acquisition software the CAMAC info needed to acquire the data. The key characters are B, C, N, and CH. The one or two character values must immediately follow the key characters, and a space should separate the fields. | KMAC | STRING | N | Y |
| --- | --- | --- | --- | --- |
| This field is meant to indicate a method to convert the data into engineering values. At the present time, the only supported value is (1), which indicates linear conversion using the form : Y = YCF1*x2 + YCF2*x + YCF3 | YCVT | SHORT | N | Y |
| Conversion Factor 1 | YCF1 | DOUBLE | N | Y |
| Conversion Factor 2 | YCF2 | DOUBLE | N | Y |
| Conversion Factor 3 | YCF3 | DOUBLE | N | Y |
| Time of First Sample which is calculated from PTS, SPRD, MNVS, and TOT. | TOFS | DOUBLE | N | N |
| This is set to “908”, set in the .dbd file. | RTYP | STRING | N | N |
| The version of the 908 record. Set in the .dbd file. | VERS | STRING | N | N |
| Same as MSTS. To link to 908 config record set. | VAL | LONG | Y | N |
| For monitors. Previous value of MNVS. | PMNV |  |  |  |
| For monitors. Previous value of MSTS. | PMST |  |  |  |
| For monitors. Previous value of EXCK. | PECK |  |  |  |
| For monitors. Previous value of PTS. | PPTS |  |  |  |
| For monitors. Previous value of RIP. | PRIP |  |  |  |
| For monitors. Previous value of TOT. | PTOT |  |  |  |
| For monitors. Previous value of TOFS. | PTOF |  |  |  |
| For monitors. Previous value of OPMD. | POPM |  |  |  |
| For monitors. Previous value of OPST. | POPS |  |  |  |
| For monitors. Previous value of SPRD | PSPR |  |  |  |
| For monitors. Previous value of SRAT. | PSRA |  |  |  |
| For monitors. Previous value of NCHN. | PNCH |  |  |  |