## Slide 1
High-k Scattering System Receiver for NSTX-U
R. Barchfeld1, E. Scott1, C.W. Domier1, P. Riemenschneider1, N.C. Luhmann Jr.1, Y. Ren2, R. Kaita2
1University of California, Davis, 2Princeton Plasma Physics Laboratory
Work supported in part by U.S. DOE Grant DE-FG02-99ER54518 and DE-AC02-09CH1146
The High-k Scattering System for NSTX-U features a 693 GHz probe beam directed from Bay G to Bay L. Scattered power is collected with receiving optics at bay L, and directed to a 4-channel subharmonic mixer array from Virginia Diodes Inc. An additional subharmonic mixer by Farran Technologies is located near the launch side of the probe beam to provide a reference channel. The probe beam and receiver are remote steerable to position the scattering volume from plasma core to edge at mid-plane ± 15 cm for various scattering schemes. This instrument will provide data on electron scale turbulence kθ from 7 to 40 cm-1, in support of anomalous transport and electron temperature gradient mode research .
693 GHz probe beam via formic acid laser
Launch Optics Bay G
Receiving Optics Bay L
NSTX-U Vacuum Vessel
Launch optics enclosure. Corrugated waveguide delivers the 693 GHz beam from above. The launch optics expand, and then focus the beam through a z-cut quartz window on Bay G. The beam waist in the center of the scattering region has a radius of ~20 mm. The beam can be steered up/down/right/left ± 2.25°.
Scattering Region
Waveguide from above
Launch Optics
Receiving optics shown in blue on bay L. Two HDPE lenses will collect and focus scattered signals. The system can translate in five axes (X, Y, Z, toroidal, and poloidal) to accommodate various measurements from core to edge at 2° to 15° scattering angles. A 4 x 1 mixer array is installed behind the second lens; however, the optics will accommodate a 8 x 1 array for later expansion in 2017.
VDI mixer array
Receiving optics
Prior to the launch optics, a beam splitter in the waveguide run will sample a small amount of power (< 1 mW) for the Farran mixer for the reference channel.
Farran Mixer
4-pixel linear array
30 mm channel spacing
x48 LO input (~14.4 GHz)
2 mm diagonal horn RF input
< 13 dB conversion loss at 693 GHz
16 dB IF amplifier gain
VDI 693 GHz 4 x 1 sub-harmonic mixer array
PCB w/
LO Source &
Phase Locking
Lo multiplier
DC bias
IF amplifiers
IF output
Digitizer
Diplexer
RF input horn
Farran Reference Mixer
LO 14.2 GHz
IF 880 MHz
Printed Circuit Board
Reference IF 880 MHz
Scattered signals RF input
Receiver Array
Reference
Box
Receiver
IF outputs
Miter Tee
HDPE Lens
FIR Out
FIR In
Example data from previous High-kr 280 GHz scattering system. Top: freq. vs time, bottom: time slices of power vs freq.
High-k Scattering Receiver